## Sheet1
|  |  | VERSION 4 |  | Hydrology Operations Manual | [5 empty] | Section No: 14.3 Appendix No: 1
|  |  | Issue Date: SEPTEMBER 2018 |  | Horizons Regional Council | [5 empty] | Portfolio: Discrete Water Quality
|  | SMARTROLL 1 | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 4 | SMARTROLL 5 | SMARTROLL 6 | SMARTROLL 7 | SMARTROLL 8 | SMARTROLL 9 | SMARTROLL 10 | SMARTROLL 11 | YSI Pro8
2016-07-01 | Date: | [7 empty] | 2016-07-05 | 2016-07-05 |  |  |  |  | Date:
| pH 7 mV Check: | [13 empty] | pH 7 mV Check:
| Electrolyte replaced Y/N: | [7 empty] | Y | Y |  |  |  |  | Electrolyte replaced Y/N:
| Post mV Check: | [13 empty] | Post mV Check:
| Depth Calibration: | [13 empty] | Depth Calibration:
| Notes: | [7 empty] | Added to fleet | Added to fleet |  |  |  |  | Notes:
|  | SMARTROLL 1 | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 4 | SMARTROLL 5 | SMARTROLL 6 | SMARTROLL 7 | SMARTROLL 8 | SMARTROLL 9
2016-07-01 | Date: | 2016-07-27 | 2016-07-27 | 2016-07-27 | 2016-07-27 | 2016-07-27 | 2016-07-27 | 2016-07-27 | 2016-07-27 | 2016-07-27 |  |  |  |  | Date:
| pH 7 mV Check: |  | OFFSITE - NO CHECKS | -20.5 | NO CHECKS- WAITING FOR NEW BARO UNIT |  |  | -15.4 | -17.5 | -8.6 |  |  |  |  | pH 7 mV Check:
| Electrolyte replaced Y/N: | Y |  | N |  | Y | Y | N | N | N |  |  |  |  | Electrolyte replaced Y/N:
| Post mV Check: | -13.1 |  |  |  | -15.6 | -13.6 | [7 empty] | Post mV Check:
| Depth Calibration: | [13 empty] | Depth Calibration:
| Notes: | [13 empty] | Notes:
|  | SMARTROLL 1 | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 4 | SMARTROLL 5 | SMARTROLL 6 | SMARTROLL 7 | SMARTROLL 8 | SMARTROLL 9
2016-08-01 | Date: |  |  |  | 2016-08-22 | 2016-08-22 | 2016-08-22 | 2016-08-22 | [6 empty] | Date:
| pH 7 mV Check: |  |  |  | BARO UNIT REPLACED sn470605 | [9 empty] | pH 7 mV Check:
| Electrolyte replaced Y/N: | [13 empty] | Electrolyte replaced Y/N:
| Post mV Check: | [13 empty] | Post mV Check:
| Depth Calibration: | [13 empty] | Depth Calibration:
| Notes: |  |  |  | RDO CAP REPLACED. | RDO CAP REPLACED | RDO CAP REPLACED | RDO CAP REPLACED | [6 empty] | Notes:
|  | SMARTROLL 1 | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 4 | SMARTROLL 5 | SMARTROLL 6 | SMARTROLL 7 | SMARTROLL 8 | SMARTROLL 9
2016-09-01 | Date: | 2016-09-05 | 2016-09-05 | 2016-09-05 | 2016-09-05 | 2016-09-05 | 2016-09-05 | 2016-09-05 | 2016-09-05 | 2016-09-05 |  |  |  |  | Date:
| pH 7 mV Check: |  | OFFSITE - NO CHECKS |  | -16.5 |  | -22.5 | -21.8 |  | -16.3 |  |  |  |  | pH 7 mV Check:
| Electrolyte replaced Y/N: | Y |  | Y | N | Y | N | N | Y | N |  |  |  |  | Electrolyte replaced Y/N:
| Post mV Check: | -11.7 |  | -10.4 |  | -14 |  |  | -12.3 | [5 empty] | Post mV Check:
| Depth Calibration: | PASS |  | PASS | PASS | PASS | PASS | PASS | FAIL | PASS |  |  |  |  | Depth Calibration:
| Notes: | [13 empty] | Notes:
|  | SMARTROLL 1 | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 4 | SMARTROLL 5 | SMARTROLL 6 | SMARTROLL 7 | SMARTROLL 8 | SMARTROLL 9
2016-09-01 | Date: | 2016-09-19 | 2016-09-19 | [11 empty] | Date:
| pH 7 mV Check: | [13 empty] | pH 7 mV Check:
| Electrolyte replaced Y/N: |  | Y | [11 empty] | Electrolyte replaced Y/N:
| Post mV Check: |  | -6.5 | [11 empty] | Post mV Check:
| Depth Calibration: | [13 empty] | Depth Calibration:
| Notes: | BARO UNIT REPLACED SN472687 | [12 empty] | Notes:
|  | SMARTROLL 1 | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 4 | SMARTROLL 5 | SMARTROLL 6 | SMARTROLL 7 | SMARTROLL 8 | SMARTROLL 9
2016-10-01 | Date: | 2016-10-03 | 2016-10-03 | 2016-10-03 | 2016-10-03 | 2016-10-03 | 2016-10-03 | 2016-10-03 | 2016-10-03 | 2016-10-03 |  |  |  |  | Date:
| pH 7 mV Check: |  |  | -11.1 | -20.1 | -10.2 |  |  | -17 | OFFSITE - NO CHECKS |  |  |  |  | pH 7 mV Check:
| Electrolyte replaced Y/N: |  | N | N | N | N | Y | Y | N | [5 empty] | Electrolyte replaced Y/N:
| Post mV Check: | [5 empty] | -10.2 | -8.2 | [6 empty] | Post mV Check:
| Depth Calibration: | FAIL | FAIL | PASS | PASS | PASS | PASS | PASS | [6 empty] | Depth Calibration:
| Notes: | SENT TO AGENT | [12 empty] | Notes:
|  | SMARTROLL 1 | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 4 | SMARTROLL 5 | SMARTROLL 6 | SMARTROLL 7 | SMARTROLL 8 | SMARTROLL 9
2016-11-01 | Date: | 2016-10-27 | 2016-10-27 | 2016-10-27 | 2016-10-27 | 2016-10-27 | 2016-10-27 | 2016-10-27 | 2016-10-27 | 2016-10-27 |  |  |  |  | Date:
| pH 7 mV Check: | SENT TO AGENT | OFFSITE - NO CHECKS | -12.9 |  | -17.9 |  |  | -19.9 | -13.2 |  |  |  |  | pH 7 mV Check:
| Electrolyte replaced Y/N: |  |  | N | Y | N | Y | Y | N | N |  |  |  |  | Electrolyte replaced Y/N:
| Post mV Check: |  |  |  | -7.1 |  | -10.6 | -18.3 | [6 empty] | Post mV Check:
| Depth Calibration: |  |  | PASS | PASS | PASS | PASS | PASS | FAIL | PASS |  |  |  |  | Depth Calibration:
| Notes: | [13 empty] | Notes:
|  | SMARTROLL 1 | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 4 | SMARTROLL 5 | SMARTROLL 6 | SMARTROLL 7 | SMARTROLL 8 | SMARTROLL 9
2016-12-01 | Date: | 2016-12-02 | 2016-12-02 | 2016-12-02 | 2016-12-02 | 2016-12-02 | 2016-12-02 | 2016-12-02 | 2016-12-02 | 2016-12-02 |  |  |  |  | Date:
| pH 7 mV Check: | SENT TO AGENT | OFFSITE - NO CHECKS | -12.9 | -11.8 | -10.2 |  |  | -19.9 | -13.2 |  |  |  |  | pH 7 mV Check:
| Electrolyte replaced Y/N: |  |  | N | N | N | Y | Y | N | N |  |  |  |  | Electrolyte replaced Y/N:
| Post mV Check: | [5 empty] | -11.4 | -18.3 | [6 empty] | Post mV Check:
| Depth Calibration: |  |  | PASS | PASS | PASS | PASS | PASS | FAIL | PASS |  |  |  |  | Depth Calibration:
| Notes: | [13 empty] | Notes:
|  | SMARTROLL 1 | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 4 | SMARTROLL 5 | SMARTROLL 6 | SMARTROLL 7 | SMARTROLL 8 | SMARTROLL 9 |  |  | YSI Pro8
2017-01-01 | Date: | 2016-12-23 | 2016-12-23 | 2016-12-23 | 2016-12-23 | 2016-12-23 | 2016-12-23 | 2016-12-23 | 2016-12-23 | 2016-12-23 |  |  | 2016-12-19 |  | Date:
| pH 7 mV Check: | SENT TO AGENT |  |  | -17.4 | -15.7 | -16.2 |  |  | -16.3 |  |  | Re-activated. New ph probe. Baro and Do checked. All passed |  | pH 7 mV Check:
| Electrolyte replaced Y/N: |  | Y | Y | N | N | N | Y | Y | N |  |  |  |  | Electrolyte replaced Y/N:
| Post mV Check: |  | -1.8 | -0.6 |  |  |  | -7.2 | -8.2 | [5 empty] | Post mV Check:
| Depth Calibration: |  | FAIL | PASS | PASS | PASS | PASS | FAIL | FAIL | PASS |  |  |  |  | Depth Calibration:
| Notes: |  |  | NEW JUNCTION |  |  |  | 1ST DEPTH FAIL | NEW JUNCTION | [5 empty] | Notes:
|  | SMARTROLL 1 | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 4 | SMARTROLL 5 | SMARTROLL 6 | SMARTROLL 7 | SMARTROLL 8 | SMARTROLL 9 |  |  | YSI Pro8
2017-02-01 | Date: | 2017-02-03 | 2017-02-03 | 2017-02-03 | 2017-02-03 | 2017-02-03 | 2017-02-03 | 2017-02-03 | 2017-02-03 | 2017-02-03 |  |  | 2017-02-03 |  | Date:
| pH 7 mV Check: |  | OFFSITE - NO CHECKS | -4.2 |  |  | -17.4 | -17 | -19.5 |  |  |  | PASSED all checks |  | pH 7 mV Check:
| Electrolyte replaced Y/N: | Y |  | N | Y | Y | N | N | N | Y |  |  |  |  | Electrolyte replaced Y/N:
| Post mV Check: | -2.2 |  |  | -7.2 | -11.9 |  |  |  | -7.2 |  |  |  |  | Post mV Check:
| Depth Calibration: | FAIL |  | PASS | PASS | PASS | PASS | FAIL | FAIL | PASS |  |  |  |  | Depth Calibration:
| Notes: | [5 empty] | NEW JUNCTION | [7 empty] | Notes:
|  | SMARTROLL 1 | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 4 | SMARTROLL 5 | SMARTROLL 6 | SMARTROLL 7 | SMARTROLL 8 | SMARTROLL 9 |  |  | YSI Pro8
2017-03-01 | Date: | 2017-03-03 | 2017-03-03 | 2017-03-03 | 2017-03-03 | 2017-03-03 | 2017-03-03 | 2017-03-03 | 2017-03-03 | 2017-03-03 |  |  | 2017-03-03 |  | Date:
| pH 7 mV Check: | -9 | -0.8 | -23 | SENT TO AGENT DUE TO CONNECTION FAILS | -20.3 | -8 | -17 | SENT TO AGENT DUE TO DEPTH FAIL | -8.3 |  |  | PASSED all checks |  | pH 7 mV Check:
| Electrolyte replaced Y/N: | N | N | N |  | N | N | N |  | N |  |  |  |  | Electrolyte replaced Y/N:
| Post mV Check: | [13 empty] | Post mV Check:
| Depth Calibration: | FAIL | PASS | PASS |  | PASS | PASS | FAIL |  | PASS |  |  |  |  | Depth Calibration:
| Notes: |  |  | NEW RDO CAP NEW PH PROBE | [10 empty] | Notes:
|  | SMARTROLL 1 | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 4 | SMARTROLL 5 | SMARTROLL 6 | SMARTROLL 7 | SMARTROLL 8 | SMARTROLL 9 |  |  | YSI Pro8
2017-04-01 | Date: | 2017-04-03 | 2017-04-03 | 2017-04-03 | 2017-04-03 | 2017-04-03 | 2017-04-03 | 2017-04-03 | 2017-04-03 | 2017-04-03 |  |  | 2017-04-03 |  | Date:
| pH 7 mV Check: | -8.5 | OFFSITE - NO CHECKS | -9.1 | SENT TO AGENT DUE TO CONNECTION FAILS | -11.5 | -17.5 |  | SENT TO AGENT DUE TO DEPTH FAIL | -14.5 |  |  | PASSED all checks |  | pH 7 mV Check:
| Electrolyte replaced Y/N: | N |  | N |  | N | N | Y |  | N |  |  |  |  | Electrolyte replaced Y/N:
| Post mV Check: | [6 empty] | -4.5 | [6 empty] | Post mV Check:
| Depth Calibration: | FAIL |  | PASS |  | PASS | PASS | FAIL |  | PASS |  |  |  |  | Depth Calibration:
| Notes: | [13 empty] | Notes:
|  | SMARTROLL 1 | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 4 | SMARTROLL 5 | SMARTROLL 6 | SMARTROLL 7 | SMARTROLL 8 | SMARTROLL 9 |  |  | YSI Pro8
2017-05-01 | Date: | 2017-05-01 | 2017-05-01 | 2017-05-01 | 2017-05-01 | 2017-05-01 | 2017-05-01 | 2017-05-01 | 2017-05-01 | 2017-05-01 |  |  | 2017-05-01 |  | Date:
| pH 7 mV Check: | -13.4 |  | -14.1 | SENT TO AGENT DUE TO CONNECTION FAILS | -16.7 | -6.3 | -9.7 | SENT TO AGENT DUE TO DEPTH FAIL | -19.7 |  |  | PASSED all checks |  | pH 7 mV Check:
| Electrolyte replaced Y/N: | N | Y | N |  | N | Y | N |  | N |  |  |  |  | Electrolyte replaced Y/N:
| Post mV Check: |  | -1.3 | [11 empty] | Post mV Check:
| Depth Calibration: | FAIL | PASS | PASS |  | PASS | PASS | FAIL |  | PASS |  |  |  |  | Depth Calibration:
| Notes: | [13 empty] | Notes:
|  | SMARTROLL 1 | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 4 | SMARTROLL 5 | SMARTROLL 6 | SMARTROLL 7 | SMARTROLL 8 | SMARTROLL 9 |  |  | YSI Pro8
2017-06-01 | Date: | 2017-06-01 | 2017-06-01 | 2017-06-01 | 2017-06-01 | 2017-06-01 | 2017-06-01 | 2017-06-01 | 2017-06-01 | 2017-06-01 |  |  | 2017-06-01 |  | Date:
| pH 7 mV Check: | -17.8 | -5.5 | -9.6 | SENT TO AGENT DUE TO CONNECTION FAILS |  | -12.3 | -14.6 | -7.5 |  |  |  | PASSED all checks |  | pH 7 mV Check:
| Electrolyte replaced Y/N: | N | N | N |  | Y | N | N | N | Y |  |  |  |  | Electrolyte replaced Y/N:
| Post mV Check: |  |  |  |  | -21.5 |  |  |  | -5.6 |  |  |  |  | Post mV Check:
| Depth Calibration: | FAIL | FAIL | PASS |  | PASS | PASS | FAIL | PASS | PASS |  |  |  |  | Depth Calibration:
| Notes: |  | 1ST FAIL | [11 empty] | Notes:
|  | SMARTROLL 1 | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 4 | SMARTROLL 5 | SMARTROLL 6 | SMARTROLL 7 | SMARTROLL 8 | SMARTROLL 9 |  |  | YSI Pro8
2017-07-01 | Date: | 2017-06-30 | 2017-06-30 | 2017-06-30 | 2017-06-30 | 2017-06-30 | 2017-06-30 | 2017-06-30 | 2017-06-30 | 2017-06-30 |  |  | 2017-06-30 |  | Date:
| pH 7 mV Check: |  | OFFSITE - NO CHECKS |  | SENT TO AGENT DUE TO CONNECTION FAILS | -14.8 | -18.6 | -6.3 | -11.5 | -14.2 |  |  | PASSED all checks |  | pH 7 mV Check:
| Electrolyte replaced Y/N: | Y |  | Y |  | N | N | N | N | N |  |  |  |  | Electrolyte replaced Y/N:
| Post mV Check: | -5.4 |  | -9.4 | [10 empty] | Post mV Check:
| Depth Calibration: | FAIL |  | PASS |  | PASS | PASS | FAIL | PASS | PASS |  |  |  |  | Depth Calibration:
| Notes: | [13 empty] | Notes:
|  | SMARTROLL 1 | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 4 | SMARTROLL 5 | SMARTROLL 6 | SMARTROLL 7 | SMARTROLL 8 | SMARTROLL 9 | SMARTROLL 10 |  | YSI Pro8
2017-08-01 | Date: | 2017-07-31 | 2017-07-31 | 2017-07-31 | 2017-07-31 | 2017-07-31 | 2017-07-31 | 2017-07-31 | 2017-07-31 | 2017-07-31 | 2017-07-31 |  | 2017-07-31 |  | Date:
| pH 7 mV Check: | -14.9 | OFFSITE - NO CHECKS | -17.2 | SENT TO AGENT DUE TO CONNECTION FAILS | -17.5 |  | SENT TO AGENT DUE TO DEPTH FAIL | -20.3 | -18.6 | UNIT ASSEMBLED VLAIDATED AND CALIBRATED. ALL PASSED ADDED TO FLEET |  | REMOVED FROM FLEET |  | pH 7 mV Check:
| Electrolyte replaced Y/N: | N |  | N |  | N | Y |  | N | N |  |  |  |  | Electrolyte replaced Y/N:
| Post mV Check: | [5 empty] | -7.14 | [7 empty] | Post mV Check:
| Depth Calibration: | FAIL |  | PASS |  | PASS | PASS |  | PASS | PASS |  |  |  |  | Depth Calibration:
| Notes: | [13 empty] | Notes:
|  | SMARTROLL 1 | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 4 | SMARTROLL 5 | SMARTROLL 6 | SMARTROLL 7 | SMARTROLL 8 | SMARTROLL 9 | SMARTROLL 10
2017-08-01 | Date: | 2017-09-19 |  |  | 2017-08-24 | 2017-08-24 | 2018-08-31 | 2017-08-24 | 2018-08-31 | 2018-08-31 |  |  |  |  | Date:
| pH 7 mV Check: | NEW RDO CAP |  |  | SENT TO AGENT DUE TO CONNECTION FAILS | NEW BARO PACK SN464265 | NEW RDO CAP | RETURNED FROM AGENT. DEPTH REPAIRED. NEW RDO MEMBRANE. PASSED CHECKS | NEW RDO CAP | NEW RDO CAP |  |  |  |  | pH 7 mV Check:
| Electrolyte replaced Y/N: | [13 empty] | Electrolyte replaced Y/N:
| Post mV Check: | [13 empty] | Post mV Check:
| Depth Calibration: | [13 empty] | Depth Calibration:
| Notes: |  |  |  |  | NEW RDO CAP | [8 empty] | Notes:
|  | SMARTROLL 1 | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 4 | SMARTROLL 5 | SMARTROLL 6 | SMARTROLL 7 | SMARTROLL 8 | SMARTROLL 9 | SMARTROLL 10
2017-10-01 | Date: | 2017-09-28 | 2017-09-28 | 2017-09-28 | 2017-09-28 | 2017-09-28 | 2017-09-28 | 2017-09-28 | 2017-09-28 | 2017-09-28 | 2017-09-28 |  |  |  | Date:
| pH 7 mV Check: |  | OFFSITE - NO CHECKS |  | SENT TO AGENT DUE TO CONNECTION FAILS |  | -16.9 | -16 |  | -23.1 |  |  |  |  | pH 7 mV Check:
| Electrolyte replaced Y/N: | Y |  | Y |  | Y | N | N | Y | N | Y |  |  |  | Electrolyte replaced Y/N:
| Post mV Check: | -11.1 |  | -11.4 |  | -17.4 |  |  | -12 |  | -7.5 |  |  |  | Post mV Check:
| Depth Calibration: | FAIL |  | PASS |  | PASS | PASS | PASS | PASS | PASS | PASS |  |  |  | Depth Calibration:
| Notes: | [13 empty] | Notes:
|  | SMARTROLL 1 | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 4 | SMARTROLL 5 | SMARTROLL 6 | SMARTROLL 7 | SMARTROLL 8 | SMARTROLL 9 | SMARTROLL 10
2017-11-01 | Date: | 2017-11-06 | 2017-11-06 | 2017-11-06 | 2017-11-06 | 2017-11-06 | 2017-11-06 | 2017-11-06 | 2017-11-06 | 2017-11-06 | 2017-11-06 |  |  |  | Date:
| pH 7 mV Check: |  | OFFSITE - NO CHECKS | -15.7 | -17.9 | -17.4 |  |  | -19.5 | OFFSITE - NO CHECKS | -14.2 |  |  |  | pH 7 mV Check:
| Electrolyte replaced Y/N: | Y |  | N | N | N | Y | Y | N |  | N |  |  |  | Electrolyte replaced Y/N:
| Post mV Check: | -9.3 |  |  |  |  | -8.7 | -8.5 | [6 empty] | Post mV Check:
| Depth Calibration: | FAIL |  | FAIL | PASS | PASS | PASS | PASS | PASS |  | PASS |  |  |  | Depth Calibration:
| Notes: |  |  | 1ST FAIL | [10 empty] | Notes:
|  | SMARTROLL 1 | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 4 | SMARTROLL 5 | SMARTROLL 6 | SMARTROLL 7 | SMARTROLL 8 | SMARTROLL 9 | SMARTROLL 10
2018-01-01 | Date: | 2017-12-22 | 2017-12-22 | 2017-12-22 | 2017-12-22 | 2017-12-22 | 2017-12-22 | 2017-12-22 | 2017-12-22 | 2017-12-22 | 2017-12-22 |  |  |  | Date:
| pH 7 mV Check: | -13 |  | -18.7 | -9 |  | -17.6 | -10.5 | -13.6 |  | -19.4 |  |  |  | pH 7 mV Check:
| Electrolyte replaced Y/N: | N | Y | N | N | Y | N | N | N | Y | N |  |  |  | Electrolyte replaced Y/N:
| Post mV Check: |  | -4.2 |  |  | -8.6 |  |  |  | -0.7 |  |  |  |  | Post mV Check:
| Depth Calibration: | FAIL | FAIL | FAIL | PASS | PASS | PASS | FAIL | PASS | PASS | PASS |  |  |  | Depth Calibration:
| Notes: | [6 empty] | 1ST FAIL (AGAIN) | [6 empty] | Notes:
|  | SMARTROLL 1 | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 4 | SMARTROLL 5 | SMARTROLL 6 | SMARTROLL 7 | SMARTROLL 8 | SMARTROLL 9 | SMARTROLL 10
2018-02-01 | Date: | 2018-02-16 | 2018-02-16 | 2018-02-16 | 2018-02-16 | 2018-02-16 | 2018-02-16 | 2018-02-16 | 2018-02-16 | 2018-02-16 | 2018-02-16 |  |  |  | Date:
| pH 7 mV Check: |  | OFFSITE - NO CHECKS |  | CONNECTION FAIL SENT TO AGENT |  |  |  | -5.6 |  | -6.1 |  |  |  | pH 7 mV Check:
| Electrolyte replaced Y/N: | Y |  | Y |  | Y | Y | Y | N | Y | N |  |  |  | Electrolyte replaced Y/N:
| Post mV Check: | -5.1 |  | -6.6 |  | -9.2 | -5.9 | -5.8 |  | -3.9 |  |  |  |  | Post mV Check:
| Depth Calibration: | FAIL |  | FAIL |  | FAIL | PASS | FAIL | PASS | PASS | PASS |  |  |  | Depth Calibration:
| Notes: |  |  |  |  | 1ST FAIL | [8 empty] | Notes:
|  | SMARTROLL 1 | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 4 | SMARTROLL 5 | SMARTROLL 6 | SMARTROLL 7 | SMARTROLL 8 | SMARTROLL 9 | SMARTROLL 10
2018-04-01 | Date: | 2018-04-13 | 2018-04-13 | 2018-04-13 | 2018-04-13 | 2018-04-13 | 2018-04-13 | 2018-04-13 | 2018-04-13 | 2018-04-13 | 2018-04-13 |  |  |  | Date:
| pH 7 mV Check: |  |  | -16.3 | CONNECTION FAIL SENT TO AGENT |  | -17.4 | -15.2 | -18.8 | -16.1 | -17.1 |  |  |  | pH 7 mV Check:
| Electrolyte replaced Y/N: | Y | Y | N |  | Y | N | N | N | N | N |  |  |  | Electrolyte replaced Y/N:
| Post mV Check: | -10.2 | NOT NOTED |  |  | -10.4 | [8 empty] | Post mV Check:
| Depth Calibration: | FAIL | FAIL | FAIL |  | FAIL | PASS | FAIL | PASS | PASS | PASS |  |  |  | Depth Calibration:
| Notes: | NEW RDO CAP | NEW RDO CAP NEW BARO UNIT |  |  |  | NEW RDO CAP |  | NEW RDO CAP | [5 empty] | Notes:
|  | SMARTROLL 1 | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 4 | SMARTROLL 5 | SMARTROLL 6 | SMARTROLL 7 | SMARTROLL 8 | SMARTROLL 9 | SMARTROLL 10
2018-05-01 | Date: | 2018-05-21 | 2018-05-21 | 2018-05-21 | 2018-05-23 | 2018-05-21 | 2018-05-21 | 2018-05-21 | 2018-05-21 | 2018-05-21 | 2018-05-21 |  |  |  | Date:
| pH 7 mV Check: | -15.7 | -8.1 |  | WRITTEN OFF. CONNECTION ISSUE UNRELIABLE. KEPT FOR SPARES | -17.1 | -11.5 | OFFSITE - NO CHECKS |  |  | -19.9 |  |  |  | pH 7 mV Check:
| Electrolyte replaced Y/N: | N | N | Y |  | N | N |  | Y | Y | N |  |  |  | Electrolyte replaced Y/N:
| Post mV Check: |  |  | -10 |  |  |  |  | -3.9 | -6.6 |  |  |  |  | Post mV Check:
| Depth Calibration: | FAIL | FAIL | FAIL |  | FAIL | PASS |  | PASS | PASS | PASS |  |  |  | Depth Calibration:
| Notes: | [7 empty] | NEW JUNCTION | [5 empty] | Notes:
|  | SMARTROLL 1 | SMARTROLL 2 | SMARTROLL 3 |  | SMARTROLL 5 | SMARTROLL 6 | SMARTROLL 7 | SMARTROLL 8 | SMARTROLL 9 | SMARTROLL 10
2018-06-01 | Date: | 2018-06-28 | 2018-06-28 | 2018-06-28 |  | 2018-06-28 | 2018-06-28 | 2018-06-28 | 2018-06-28 | 2018-06-28 | 2018-06-28 |  |  |  | Date:
| pH 7 mV Check: | [5 empty] | -20.9 |  | -11.1 | [5 empty] | pH 7 mV Check:
| Electrolyte replaced Y/N: | Y | Y | Y |  | Y | N | Y | N | Y | Y |  |  |  | Electrolyte replaced Y/N:
| Post mV Check: | -10 | -10.5 | -14.2 |  | -13.2 |  | -7.9 |  | -6 | -9.9 |  |  |  | Post mV Check:
| Depth Calibration: | FAIL | FAIL | FAIL |  | FAIL | PASS | FAIL | PASS | PASS | PASS |  |  |  | Depth Calibration:
| Notes: |  | NEW PH PROBE | [5 empty] | NEW PH PROBE | NEW PH PROBE | DEPTH CAL INITALLY FAILED! |  |  |  | Notes:
|  | SMARTROLL 1 | SMARTROLL 2 | SMARTROLL 3 |  | SMARTROLL 5 | SMARTROLL 6 | SMARTROLL 7 | SMARTROLL 8 | SMARTROLL 9 | SMARTROLL 10 | SMARTROLL 11
2018-07-01 | Date: | 2018-07-30 | 2018-07-30 | 2018-07-30 |  | 2018-07-30 | 2018-07-30 | 2018-07-30 | 2018-07-30 | 2018-07-30 | 2018-07-30 | 2018-08-10 |  |  | Date:
| pH 7 mV Check: | -14 | -17.3 | -15.6 |  |  |  | OFFSITE - NO CHECKS | -18.4 | -14.9 | -8 | UNIT ASSEMBLED VLAIDATED AND CALIBRATED. ALL PASSED ADDED TO FLEET |  |  | pH 7 mV Check:
| Electrolyte replaced Y/N: | N | N | N |  | Y | Y |  | N | N | N |  |  |  | Electrolyte replaced Y/N:
| Post mV Check: |  |  |  |  | -15.4 | -11.2 | [7 empty] | Post mV Check:
| Depth Calibration: | FAIL | FAIL | FAIL |  | FAIL | PASS |  | PASS | PASS | PASS |  |  |  | Depth Calibration:
| Notes: | [9 empty] | NEW PH PROBE 27/07/18 |  |  |  | Notes:
|  | SMARTROLL 1 | SMARTROLL 2 | SMARTROLL 3 |  | SMARTROLL 5 | SMARTROLL 6 | SMARTROLL 7 | SMARTROLL 8 | SMARTROLL 9 | SMARTROLL 10 | SMARTROLL 11
2018-08-01 | Date: | 2018-08-31 | 2018-08-31 | 2018-08-31 |  | 2018-08-31 | 2018-08-31 | 2018-08-31 | 2018-08-31 | 2018-08-31 | 2018-08-31 | 2018-08-31 |  |  | Date:
| pH 7 mV Check: | -23 | -22.7 | -24 |  | -25 | -24.6 | -15.5 | -25.2 | -19.9 | -13.4 | -10.2 |  |  | pH 7 mV Check:
| Electrolyte replaced Y/N: | Y | Y | N |  | Y | Y | N | Y | N | N | N |  |  | Electrolyte replaced Y/N:
| Post mV Check: | -9 | -6 | N/A |  | -10 | -5 | N/A | -14.6 | N/A | N/A | N/A |  |  | Post mV Check:
| Depth Calibration: | FAIL | FAIL | FAIL |  | FAIL | PASS | FAIL | PASS | PASS | PASS | PASS |  |  | Depth Calibration:
| RDO Days: | 295 | 280 | 379 |  | 84 | 295 | 84 | 295 | 91 | 29 | 435
| Notes: | [13 empty] | Notes:
|  | SMARTROLL 1 | SMARTROLL 2 | SMARTROLL 3 |  | SMARTROLL 5 | SMARTROLL 6 | SMARTROLL 7 | SMARTROLL 8 | SMARTROLL 9 | SMARTROLL 10 | SMARTROLL 11
2018-09-01 | Date: | 2018-09-27 | 2018-09-27 | 2018-09-27 |  | 2018-09-27 | 2018-09-27 | 2018-09-27 | 2018-09-27 | 2018-09-27 | 2018-09-27 | 2018-09-27 |  |  | Date:
| pH 7 mV Check: | -13.7 | OFFSITE - NO CHECKS | -24.7 |  | -19.2 | -13.1 | -17 | -13 | 2-22.7 | 17.7 | -12.3 |  |  | pH 7 mV Check:
| Electrolyte replaced Y/N: | N |  | Y |  | N | N | N | N | Y | New RDO Cap | N |  |  | Electrolyte replaced Y/N:
| Post mV Check: | N/A |  | -5.6 |  | N/A | N/A | N/A | N/A | -4 | N/A | N/A |  |  | Post mV Check:
| Depth Calibration: | FAIL |  | FAIL |  | FAIL | PASS | FAIL | PASS | PASS | PASS | PASS |  |  | Depth Calibration:
| RDO Days: | 268 |  | 352 |  | 456 | 268 | 456 | 268 | 456 | 456 | 408
| Notes: |  |  |  |  | New RDO Cap |  | New RDO Cap |  | New RDO Cap | New RDO Cap |  |  |  | Notes:
|  | SMARTROLL 1 | SMARTROLL 2 | SMARTROLL 3 |  | SMARTROLL 5 | SMARTROLL 6 | SMARTROLL 7 | SMARTROLL 8 | SMARTROLL 9 | SMARTROLL 10 | SMARTROLL 11
2018-10-01 | Date: | 2018-11-01 | 2018-11-01 | 2018-11-01 |  | 2018-11-01 | 2018-11-01 | 2018-11-01 | 2018-11-01 | 2018-11-01 | 2018-11-01 | 2018-11-01 |  |  | Date:
| pH 7 mV Check: | WRITTEN OFF | OFFSITE - NO CHECKS | -23 |  | -35.2 | -35.4 | -25.3 | -20.8 | -16.5 | -28.8 | -20.8 |  |  | pH 7 mV Check:
| Electrolyte replaced Y/N: |  |  | Y |  | Y | Y | Y | Y | New RDO Cap | Y | Y |  |  | Electrolyte replaced Y/N:
| Post mV Check: |  |  | -9.6 |  | -11.4 | -8.8 | -4.7 | -10.9 | N/A | -8.1 | -11.6 |  |  | Post mV Check:
| Depth Calibration: |  |  | FAIL |  | FAIL | PASS | FAIL | PASS | PASS | PASS | PASS |  |  | Depth Calibration:
| RDO Days: |  |  | 317 |  | 421 | 233 | 421 | 233 | 421 | 421 | 373
| Notes: |  |  |  |  | pH slow to respond | pH slow to respond | [7 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 5 | SMARTROLL 6 | SMARTROLL 7 | SMARTROLL 8 | SMARTROLL 9 | SMARTROLL 10 | SMARTROLL 11 | SMARTROLL 12 | SMARTROLL 13
2018-11-01 | Date: | 2018-11-30 | 2018-11-30 | 2018-11-30 | 2018-12-03 | 2018-11-30 | 2018-11-30 | 2018-11-30 | 2018-12-03 | 2018-11-30 | 2018-11-29 | 2018-11-29 |  |  | Date:
| pH 7 mV Check: | OFFSITE - NO CHECKS | -17 | -19 | -35.7 | -12.8 | -17.3 | -22 | -22.5 | -15.1 | -2 | -3 |  |  | pH 7 mV Check:
| Electrolyte replaced Y/N: |  | N | N | Y | N | N | Y | N | N | N | N |  |  | Electrolyte replaced Y/N:
| Post mV Check: |  | N/A | N/A | -10.2 | N/A | N/A | -6.6 | N/A | N/A | N/A | N/A |  |  | Post mV Check:
| Depth Calibration: |  | FAIL | FAIL | PASS | FAIL | PASS | PASS | PASS | PASS | FAIL | PASS |  |  | Depth Calibration:
| RDO Days: |  | 288 | 392 | 201 | 392 | 204 | 392 | 389 | 344 | 456 | 456 |  |  | RDO Days:
| Notes: | [13 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 5 | SMARTROLL 6 | SMARTROLL 7 | SMARTROLL 8 | SMARTROLL 9 | SMARTROLL 10 | SMARTROLL 11 | SMARTROLL 12 | SMARTROLL 13
2018-12-01 | Date: | 2018-12-21 | 2018-12-21 | 2018-12-21 | 2018-12-21 | 2018-12-21 | 2018-12-21 | 2018-12-21 | 2018-12-21 | 2018-12-21 | 2018-12-21 | 2018-12-21 |  |  | Date:
| pH 7 mV Check: | OFFSITE - NO CHECKS | -19.3 | -17.7 | -25.6 | -17.2 | -17.9 | -11.3 | -13.7 | -14.2 | -8.6 | -9.8 |  |  | pH 7 mV Check:
| Electrolyte replaced Y/N: |  | N | N | Y | N | N | N | N | N | N | N |  |  | Electrolyte replaced Y/N:
| Post mV Check: |  | N/A | N/A | -42.8 | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  |  | Post mV Check:
| Depth Calibration: |  | FAIL | PASS | PASS | FAIL | PASS | PASS | PASS | PASS | FAIL | PASS |  |  | Depth Calibration:
| RDO Days: |  | 267 | 371 | 183 | 371 | 183 | 371 | 371 | 323 | 434 | 434 |  |  | RDO Days:
| Notes: |  |  |  | probe connection within sonde damaged | [9 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 5 | SMARTROLL 6 | SMARTROLL 7 | SMARTROLL 8 | SMARTROLL 9 | SMARTROLL 10 | SMARTROLL 11 | SMARTROLL 12 | SMARTROLL 13
2019-01-01 | Date: | [13 empty] | Date:
| pH 7 mV Check: | [13 empty] | pH 7 mV Check:
| Electrolyte replaced Y/N: | [13 empty] | Electrolyte replaced Y/N:
| Post mV Check: | [13 empty] | Post mV Check:
| Depth Calibration: | [13 empty] | Depth Calibration:
| RDO Days: | [13 empty] | RDO Days:
| Notes: | [13 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 5 | SMARTROLL 6 | SMARTROLL 7 | SMARTROLL 8 | SMARTROLL 9 | SMARTROLL 10 | SMARTROLL 11 | SMARTROLL 12 | SMARTROLL 13
2019-02-01 | Date: | 2019-02-18 | 2019-02-01 | 2019-02-01 | WRITTEN OFF 07/01/19 | 2019-02-01 | 2019-02-01 | 2019-02-01 | 2019-02-01 | 2019-02-01 | 2019-02-01 | 2019-02-01 |  |  | Date:
| pH 7 mV Check: | -44.2 | -31.6 | -21.6 |  | -20.1 | -23.6 | -18 | -29.8 | -21.7 | -19 | -19.2 |  |  | pH 7 mV Check:
| Electrolyte replaced Y/N: | Y | Y | Y |  | Y | Y | N | N | Y | N | Y |  |  | Electrolyte replaced Y/N:
| Post mV Check: | -1.4 | -6.8 | -4.4 |  | -3.5 | -7 | N/A | -7.7 | -7.2 | N/A | -4.4 |  |  | Post mV Check:
| Depth Calibration: | not checked | FAIL | FAIL |  | FAIL | PASS | PASS | PASS | PASS | FAIL | PASS |  |  | Depth Calibration:
| RDO Days: | not checked | 225 | 329 |  | 329 | 141 | 329 | 329 | 281 | 392 | 439 |  |  | RDO Days:
| Notes: | OFFSITE - on 1st feb | [9 empty] | returned from thermo 1190131 |  |  | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 5 | SMARTROLL 6 | SMARTROLL 7 | SMARTROLL 8 | SMARTROLL 9 | SMARTROLL 10 | SMARTROLL 11 | SMARTROLL 12 | SMARTROLL 13
2019-03-01 | Date: | 2019-03-01 | 2019-03-01 | 2019-03-01 | WRITTEN OFF 07/01/19 | 2019-03-01 | 2019-03-01 | 2019-03-01 | 2019-03-01 | 2019-03-01 | 2019-03-01 | 2019-03-01 |  |  | Date:
| pH 7 mV Check: | OFFSITE - NO CHECKS | -17.4 | -8.9 |  | -12.6 | -14.6 | -18 | MISSING | -14.5 | -27.5 | -9 |  |  | pH 7 mV Check:
| Electrolyte replaced Y/N: |  | N | N |  | N | N | N |  | Y | Y | N |  |  | Electrolyte replaced Y/N:
| Post mV Check: |  | N/A | N/A |  | N/A | N/A | N/A |  | N/A | -5.2 | N/A |  |  | Post mV Check:
| Depth Calibration: |  | FAIL | FAIL |  | FAIL | PASS | PASS |  | PASS | FAIL | PASS |  |  | Depth Calibration:
| RDO Days: |  | 197 | 301 |  | 301 | 113 | 301 |  | 253 | 364 | 411 |  |  | RDO Days:
| Notes: | [13 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 5 | SMARTROLL 6 | SMARTROLL 7 | SMARTROLL 8 | SMARTROLL 9 | SMARTROLL 10 | SMARTROLL 11 | SMARTROLL 12 | SMARTROLL 13
2019-04-01 | Date: | 2019-05-03 | 2019-05-03 | 2019-05-03 | WRITTEN OFF 07/01/19 | 2019-05-03 | DAMAGED IN THE FIELD ON 14 MARCH 2018 & SUBSEQUENTLY WRITTEN OFF | 2019-05-03 | 2019-05-03 | 2019-05-03 | 2019-05-03 | 2019-05-03 |  |  | Date:
| pH 7 mV Check: | -29.2 | -19.7 | -20 |  | -7.2 |  | -16.4 | -33.2 | -24.7 | -22.7 | -27.5 |  |  | pH 7 mV Check:
| Electrolyte replaced Y/N: | Y | Y | Y |  | N |  | N | Y | Y | Y | Y |  |  | Electrolyte replaced Y/N:
| Post mV Check: | -6.6 | -14.8 | -14.1 |  | N/A |  | N/A | -11.9 | -12.9 | -8.4 | -8.4 |  |  | Post mV Check:
| Depth Calibration: | FAIL | FAIL | FAIL |  | FAIL |  | PASS | PASS | PASS | FAIL | PASS |  |  | Depth Calibration:
| RDO Days: | 35 | 134 | 238 |  | 238 |  | 238 | 238 | 190 | 301 | 348 |  |  | RDO Days:
| Notes: | [13 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 5 | SMARTROLL 6 | SMARTROLL 7 | SMARTROLL 8 | SMARTROLL 9 | SMARTROLL 10 | SMARTROLL 11 | SMARTROLL 12 | SMARTROLL 13
2019-05-01 | Date: | 2019-05-24 | 2019-05-24 | 2019-05-24 | WRITTEN OFF 07/01/19 | 2019-05-24 | DAMAGED IN THE FIELD ON 14 MARCH 2018 & SUBSEQUENTLY WRITTEN OFF | 2019-05-24 | 2019-05-24 | 2019-05-24 | 2019-05-24 | 2019-05-24 |  |  | Date:
| pH 7 mV Check: | -9.9 | -21.2 | -13.3 |  | -15.9 |  | -16.6 | -15.8 | -22.3 | -10.9 | -11.6 |  |  | pH 7 mV Check:
| Electrolyte replaced Y/N: | N | Y | N |  | N |  | N | N | Y | N | N |  |  | Electrolyte replaced Y/N:
| Post mV Check: | N/A | -11.8 | N/A |  | N/A |  | N/A | N/A | -13.2 | N/A | N/A |  |  | Post mV Check:
| Depth Calibration: | FAIL | FAIL | FAIL |  | FAIL |  | PASS | PASS | PASS | FAIL | PASS |  |  | Depth Calibration:
| RDO Days: | 456 | 113 | 217 |  | 217 |  | 217 | 217 | 169 | 280 | 327 |  |  | RDO Days:
| Notes: | RDO Cap replaced |  | PH/ORP Probe replaced | [10 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 5 | SMARTROLL 6 | SMARTROLL 7 | SMARTROLL 8 | SMARTROLL 9 | SMARTROLL 10 | SMARTROLL 11 | SMARTROLL 12 | SMARTROLL 13
2019-06-01 | Date: | 2019-07-03 | 2019-07-03 | 2019-07-03 | WRITTEN OFF 07/01/19 | 2019-07-03 | DAMAGED IN THE FIELD ON 14 MARCH 2018 & SUBSEQUENTLY WRITTEN OFF | 2019-07-03 | 2019-07-03 | 2019-07-03 | 2019-07-03 | 2019-07-03 |  |  | Date:
| pH 7 mV Check: | NOT CHECKED | -19.1 | -21.8 |  | -20.3 |  | -26.5 | -23.3 | -17.4 | -20.8 | -16.1 |  |  | pH 7 mV Check:
| Electrolyte replaced Y/N: |  | N | Y |  | Y |  | Y | Y | N | Y | N |  |  | Electrolyte replaced Y/N:
| Post mV Check: |  | N/A | -6.3 |  | -7.4 |  | -7.6 | -9.4 | N/A | -8.1 | N/A |  |  | Post mV Check:
| Depth Calibration: |  | FAIL | FAIL |  | FAIL |  | PASS | PASS | PASS | FAIL | PASS |  |  | Depth Calibration:
| RDO Days: |  | 73 | 177 |  | 177 |  | 177 | 177 | 106 | 257 | 287 |  |  | RDO Days:
| Notes: | [13 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 5 | SMARTROLL 6 | SMARTROLL 7 | SMARTROLL 8 | SMARTROLL 9 | SMARTROLL 10 | SMARTROLL 11 | SMARTROLL 12 | SMARTROLL 13
JUL/AUG 2019 | Date: | 2019-07-26 | 2019-07-26 | 2019-07-26 | WRITTEN OFF 07/01/19 | 2019-07-26 | DAMAGED IN THE FIELD ON 14 MARCH 2018 & SUBSEQUENTLY WRITTEN OFF | 2019-07-26 | 2019-07-26 | 2019-07-26 | 2019-07-26 | 2019-05-24 |  |  | Date:
| pH 7 mV Check: | -21.6 | -20.4 | -11 |  | -10.8 |  | -12.8 | -14 | -19.4 | -0.7 | -19.1 |  |  | pH 7 mV Check:
| Electrolyte replaced Y/N: | Y | N | N |  | N |  | N | N | N | see notes | N |  |  | Electrolyte replaced Y/N:
| Post mV Check: | -3.1 | N/A | N/A |  | N/A |  | N/A | N/A | N/A | see notes | N/A |  |  | Post mV Check:
| Depth Calibration: | FAIL | FAIL | FAIL |  | FAIL |  | PASS | PASS | PASS | FAIL | PASS |  |  | Depth Calibration:
| RDO Days: | 393 | 50 | 154 |  | 154 |  | 154 | 154 | 106 | 280 | 264 |  |  | RDO Days:
| Notes: | [9 empty] | PH/ORP Probe replaced |  |  |  | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 5 | SMARTROLL 6 | SMARTROLL 7 | SMARTROLL 8 | SMARTROLL 9 | SMARTROLL 10 | SMARTROLL 11 | SMARTROLL 12 | SMARTROLL 13
2019-09-01 | Date: | 2019-09-05 | 2019-09-05 | 2019-09-05 | WRITTEN OFF 07/01/19 | 2019-09-05 | DAMAGED IN THE FIELD ON 14 MARCH 2018 & SUBSEQUENTLY WRITTEN OFF | 2019-09-23 | 2019-09-05 | 2019-09-05 | 2019-09-05 | 2019-09-23 |  |  | Date:
| pH 7 mV Check: | -8.4 | -54 | -15.7 |  | -19.9 |  | -8.6 | -13 | -27.9 | -8.4 | -27.5 |  |  | pH 7 mV Check:
| Electrolyte replaced Y/N: | N | Y | N |  | N |  | Y | N | Y | N | Y |  |  | Electrolyte replaced Y/N:
| Post mV Check: | N/A | -10.4 | N/A |  | N/A |  | above value | N/A | -10.6 | N/A | -13.5 |  |  | Post mV Check:
| Depth Calibration: | FAIL | FAIL | FAIL |  | FAIL |  | PASS | PASS | PASS | FAIL | PASS |  |  | Depth Calibration:
| RDO Days: |  | 429 | 113 |  | 113 |  | 95 | 113 | 429 | 176 | 223 |  |  | RDO Days:
| Notes: |  | RDO cap replaced 9/9/19 |  |  |  |  | PH/ORP Probe replaced |  | RDO cap replaced 9/9/19 |  |  |  |  | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 5 | SMARTROLL 7 | SMARTROLL 9 | SMARTROLL 10 | SMARTROLL 11 | SMARTROLL 12 | SMARTROLL 13 | AQUATROLL 1
2019-10-01 | Date: | 2019-10-07 | 2019-10-07 | 2019-10-07 | 2019-10-07 | 2019-10-07 | 2019-10-07 | 2019-10-07 | 2019-10-07 | 2019-10-07 | OFFSITE - WITH MASSEY DRONE PROJECT |  |  |  | Date:
| pH 7 mV Check: | -17.5 | -13.6 | -13.6 | -23.6 | -9.3 | -13.6 | -10.3 | -12.2 | -14.7 |  |  |  |  | pH 7 mV Check:
| Electrolyte replaced Y/N: | N | N | N | y | N | N | N | N | N |  |  |  |  | Electrolyte replaced Y/N:
| Post mV Check: | N/A | N/A | N/A | -9.8 | N/A | N/A | N/A | N/A | N/A |  |  |  |  | Post mV Check:
| Depth Calibration: | FAIL | FAIL | FAIL | FAIL | PASS | PASS | PASS | FAIL | PASS |  |  |  |  | Depth Calibration:
| RDO Days: | 320 | 397 | 81 | 81 | 81 | 81 | 397 | 397 | 191 |  |  |  |  | RDO Days:
| Notes: | [13 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 5 | SMARTROLL 7 | SMARTROLL 9 | SMARTROLL 10 | SMARTROLL 11 | SMARTROLL 12 | SMARTROLL 13 | AQUATROLL 1
2019-11-01 | Date: | 2019-11-04 | 2019-11-04 | 2019-11-04 | 2019-11-04 | 2019-11-04 | 2019-11-04 | 2019-11-04 | 2019-11-04 | 2019-11-04 | OFFSITE - WITH MASSEY DRONE PROJECT |  |  |  | Date:
| pH 7 mV Check: | Not checked | -2.3 | -23.4 | -12.9 | 5.4 | -5.5 | -14 | -18.2 | -21.4 |  |  |  |  | pH 7 mV Check:
| Electrolyte replaced Y/N: |  | Y | Y | N | Y | Y | N | N | Y |  |  |  |  | Electrolyte replaced Y/N:
| Post mV Check: |  | post probe install | 0 | N/A | post probe install | post probe install | N/A | N/A | -4.5 |  |  |  |  | Post mV Check:
| Depth Calibration: |  | FAIL | FAIL | FAIL | PASS | PASS | PASS | FAIL | PASS |  |  |  |  | Depth Calibration:
| RDO Days: |  | 369 | 53 | 53 | 53 | 53 | 369 | 369 | 163 |  |  |  |  | RDO Days:
| Notes: |  | New pH/ORP probe installed |  |  | New pH/ORP probe installed | New pH/ORP probe installed | [7 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 5 | SMARTROLL 7 | SMARTROLL 9 | SMARTROLL 10 | SMARTROLL 11 | SMARTROLL 12 | SMARTROLL 13 | AQUATROLL 1
2019-12-01 | Date: | 2019-11-29 | 2019-11-29 | 2019-11-29 | 2019-11-29 | 2019-11-29 | 2019-11-29 | 2019-11-29 | 2019-11-29 | 2019-11-29 | Not checked |  |  |  | Date:
| pH 7 mV Check: | Not on site | -6.4 | -6.3 | -15.9 | -0.8 | -0.2 | -19.3 | -24.9 | NOT checked - waiting for Baro unit to be fixed |  |  |  |  | pH 7 mV Check:
| Electrolyte replaced Y/N: |  | N | N | N | N | N | N | N | [5 empty] | Electrolyte replaced Y/N:
| Post mV Check: |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | [5 empty] | Post mV Check:
| Depth Calibration: |  | FAIL | FAIL | FAIL | PASS | PASS | PASS | FAIL | [5 empty] | Depth Calibration:
| RDO Days: |  | 344 | 456 | 456 | 456 | 456 | 344 | 91 | [5 empty] | RDO Days:
| Notes: |  |  | RDO cap replaced | RDO cap replaced | RDO cap replaced | RDO cap replaced | [7 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 5 | SMARTROLL 7 | SMARTROLL 9 | SMARTROLL 10 | SMARTROLL 11 | SMARTROLL 12 | SMARTROLL 13 | AQUATROLL 1
2019-12-01 | Date: | 2019-12-02 |  |  | 2019-12-06 |  |  | 2019-12-06 |  |  | Not checked |  |  |  | Date:
| pH 7 mV Check: | -34.5 |  |  | -32.4 |  |  | -4.9 | [6 empty] | pH 7 mV Check:
| Electrolyte replaced Y/N: | Y |  |  | Y |  |  | Y | [6 empty] | Electrolyte replaced Y/N:
| Post mV Check: | -8.4 |  |  | -6.1 |  |  | n/a | [6 empty] | Post mV Check:
| Depth Calibration: | FAIL |  |  | FAIL |  |  | PASS | [6 empty] | Depth Calibration:
| RDO Days: | 264 |  |  | 337 |  |  | 337 | [6 empty] | RDO Days:
| Notes: | [6 empty] | New pH/ORP probe installed | [6 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 5 | SMARTROLL 7 | SMARTROLL 9 | SMARTROLL 10 | SMARTROLL 11 | SMARTROLL 12 | SMARTROLL 13 | AQUATROLL 1
2019-12-01 | Date: | 2019-12-19 | 2019-12-19 | 2019-12-19 | 2019-12-19 | 2019-12-19 | 2019-12-19 | 2019-12-19 | 2019-12-19 | 2019-12-19 | 2019-12-19 |  |  |  | Date:
| pH 7 mV Check: | -11.3 | -20.1 | -22.9 | -11.7 | -13.2 | -16.3 | -9.9 | -13.3 | -15.8 | Not checked |  |  |  | pH 7 mV Check:
| Electrolyte replaced Y/N: | N | N | N | N | N | N | N | N | N |  |  |  |  | Electrolyte replaced Y/N:
| Post mV Check: | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  |  |  |  | Post mV Check:
| Depth Calibration: | FAIL | FAIL | FAIL | FAIL | PASS | PASS | PASS | FAIL | PASS |  |  |  |  | Depth Calibration:
| RDO Days: | 247 | 324 | 436 | 436 | 436 | 436 | 324 | 71 | 118 |  |  |  |  | RDO Days:
| Notes: | stored in pH4 buffer for closed | stored in pH4 buffer for closed | stored in pH4 buffer for closed | stored in pH4 buffer for closed | stored in pH4 buffer for closed | stored in pH4 buffer for closed | stored in pH4 buffer for closed | stored in pH4 buffer for closed | stored in pH4 buffer for closed | stored in pH4 buffer for closed |  |  |  | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 5 | SMARTROLL 7 | SMARTROLL 9 | SMARTROLL 10 | SMARTROLL 11 | SMARTROLL 12 | SMARTROLL 13 | AQUATROLL 1
JAN/FEB 2020 | Date: | 2020-02-04 | 2020-02-04 | 2020-02-04 | 2020-02-04 | 2020-02-04 | 2020-02-04 | 2020-02-04 | 2020-02-04 | 2020-02-04 | 2019-12-19 |  |  |  | Date:
| pH 7 mV Check: | -25.9 | -28.6 | -31.4 | -19.3 | -26.1 | -28 | -19.7 | -24.1 | -15.8 | Not checked |  |  |  | pH 7 mV Check:
| Electrolyte replaced Y/N: | Y | Y | Y | N | Y | Y | N | Y | N |  |  |  |  | Electrolyte replaced Y/N:
| Post mV Check: | -6.9 | -5.2 | -9.6 | N/A | -3.7 | -4.8 | N/A | -8.2 | N/A |  |  |  |  | Post mV Check:
| Depth Calibration: | FAIL | FAIL | FAIL | FAIL | PASS | PASS | PASS | FAIL | PASS |  |  |  |  | Depth Calibration:
| RDO Days: | 200 | 277 | 389 | 389 | 298 | 389 | 185 | 456 | 338 |  |  |  |  | RDO Days:
| Notes: | [13 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 5 | SMARTROLL 7 | SMARTROLL 9 | SMARTROLL 10 | SMARTROLL 11 | SMARTROLL 12 | SMARTROLL 13 | AQUATROLL 1
Feb/Mar 2020 | Date: | 2020-02-28 | 2020-02-28 | 2020-02-28 | 2020-02-28 | 2020-02-28 | 2020-02-28 | 2020-02-28 | 2020-02-28 | 2020-02-28 | 2020-02-28 |  |  |  | Date:
| pH 7 mV Check: | Not on site | -16.2 | -25.1 | -29.5 | -11.3 | -14.9 | -27.5 | -17.3 | -27.5 | -37.8 |  |  |  | pH 7 mV Check:
| Electrolyte replaced Y/N: |  | N | Y | Y | N | N | Y | N | Y | Y |  |  |  | Electrolyte replaced Y/N:
| Post mV Check: |  | N/A | -10.4 | -9.7 | N/A | N/A | -6.5 | N/A | -8.6 | -1.5 |  |  |  | Post mV Check:
| Depth Calibration: |  | FAIL | FAIL | FAIL | PASS | PASS | PASS | FAIL | PASS | PASS |  |  |  | Depth Calibration:
| RDO Days: |  | 263 | 365 | 365 | 274 | 365 | 162 | 432 | 314 | ?? |  |  |  | RDO Days:
| Notes: | [13 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 5 | SMARTROLL 7 | SMARTROLL 9 | SMARTROLL 10 | SMARTROLL 11 | SMARTROLL 12 | SMARTROLL 13 | AQUATROLL 1
2020-05-01 | Date: | 2020-05-06 | 2020-05-06 | 2020-05-06 | 2020-05-06 | 2020-05-06 | 2020-05-06 | 2020-05-06 | 2020-05-06 | 2020-05-06 | 2020-05-06 |  |  |  | Date:
| pH 7 mV Check: | Not on site | -25.2 | -22.3 | Not detecting pH/ORP sensor - taken out of fleet | -28.1 | In Marton - Unable to check | -17.9 | -30.1 | -24.3 | -19.2 |  |  |  | pH 7 mV Check:
| Electrolyte replaced Y/N: |  | Y | Y |  | N |  | Y | N | Y | Y |  |  |  | Electrolyte replaced Y/N:
| Post mV Check: |  | -6.2 | -12.8 |  | -6.6 |  | -3.2 | -10.3 | -7.5 | -2.5 |  |  |  | Post mV Check:
| Depth Calibration: |  | FAIL | FAIL |  | PASS |  | PASS | FAIL | PASS | PASS |  |  |  | Depth Calibration:
| RDO Days: |  | 185 | 297 |  | 206 |  | 93 | 364 | 247 | ?? |  |  |  | RDO Days:
| Notes: |  |  |  | No checks | [5 empty] | RDO Cal PASS |  |  |  | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 5 | SMARTROLL 7 | SMARTROLL 9 | SMARTROLL 10 | SMARTROLL 11 | SMARTROLL 12 | SMARTROLL 13 | AQUATROLL 1
2020-06-03 | Date: | 2020-06-03 | 2020-06-03 | 2020-06-03 | 2020-06-03 | 2020-06-03 | 2020-06-03 | 2020-06-04 | 2020-06-03 | 2020-06-03 | 2020-06-04 |  |  |  | Date:
| pH 7 mV Check: | Not on site | -14.8 | -18.5 | Broken. Unoperable. | -8.2 | -34.2 | -12.3 | -17.3 | -16.1 | -8.8 |  |  |  | pH 7 mV Check:
| Electrolyte replaced Y/N: |  | N | N |  | N | Y | N | N | N | N |  |  |  | Electrolyte replaced Y/N:
| Post mV Check: |  | N/A | N/A |  | N/A | -4.5 | N/A | N/A | N/A | N/A |  |  |  | Post mV Check:
| Depth Calibration: |  | FAIL | FAIL |  | PASS | PASS | PASS | FAIL | PASS | PASS |  |  |  | Depth Calibration:
| RDO Days: |  | 157 | 269 |  | 178 | 269 | 65 | 336 | 219 | ?? |  |  |  | RDO Days:
| Notes: |  |  |  | No checks |  | Batteries changed. | Checked day after others. |  | Batteries changed. | Checked day after others. |  |  |  | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 5 | SMARTROLL 7 | SMARTROLL 9 | SMARTROLL 10 | SMARTROLL 11 | SMARTROLL 12 | SMARTROLL 13 | AQUATROLL 1
2020-07-02 | Date: | 2020-07-02 | 2020-07-02 | 2020-07-02 |  | 2020-07-02 | 2020-07-02 | 2020-07-02 | 2020-07-02 | 2020-07-02 | 2020-07-02 |  |  |  | Date:
| pH 7 mV Check: | -33 | -23.4 | -25.2 | Broken. Unoperable. | -12.1 | -16.5 | -20.7 | -23.7 | -25.6 | -10.9 |  |  |  | pH 7 mV Check:
| Electrolyte replaced Y/N: | Y | Y | Y |  | N | N | Y | Y | Y | N |  |  |  | Electrolyte replaced Y/N:
| Post mV Check: | -6.2 | -9.1 | -16.2 |  | N/A | N/A | -8.3 | -11.1 | -10.8 | N/A |  |  |  | Post mV Check:
| Depth Calibration: | FAIL | FAIL | FAIL |  | PASS | PASS | PASS | FAIL | PASS | PASS |  |  |  | Depth Calibration:
| RDO Days: | 422 | 128 | 240 |  | 149 | 240 | 37 | 307 | 190 | ?? |  |  |  | RDO Days:
| Notes: | [13 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 5 | SMARTROLL 7 | SMARTROLL 9 | SMARTROLL 10 | SMARTROLL 11 | SMARTROLL 12 | SMARTROLL 13 | AQUATROLL 1 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4
2020-08-04 | Date: | 2020-08-04 | 2020-08-04 | 2020-08-04 | 2020-08-04 | 2020-08-04 | 2020-08-04 | 2020-08-04 | 2020-08-04 | 2020-08-04 | 2020-08-04 | 2020-08-04 | 2020-08-04 | 2020-08-04 | Date:
| pH 7 mV Check: | Not on site | -16.3 | -9.5 | Broken. Unoperable. | -19.1 | -25.2 | -16.1 | -18.3 | -19.8 | -18.9 | -11.9 | -5.6 | -12.8 | pH 7 mV Check:
| Electrolyte replaced Y/N: |  | N | N |  | Y | Y | N | Y | Y | Y | N | N | N | Electrolyte replaced Y/N:
| Post mV Check: |  | N/A | N/A |  | -2.3 | -0.6 | N/A | -7.8 | -1.7 | 1.7 | N/A | N/A | N/A | Post mV Check:
| Depth Calibration: |  | FAIL | FAIL |  | PASS | PASS | PASS | FAIL | FAIL | PASS | PASS | PASS | PASS | Depth Calibration:
| RDO Days: |  | 95 | 207 |  | 116 | 207 | 352 | 274 | 157 | ?? | ?? | ?? | ?? | RDO Days:
| Notes: | [13 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 5 | SMARTROLL 7 | SMARTROLL 9 | SMARTROLL 10 | SMARTROLL 11 | SMARTROLL 12 | SMARTROLL 13 | AQUATROLL 1 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4
2020-08-27 | Date: | 2020-08-27 |  |  |  | 2020-08-27 | [8 empty] | Date:
| pH 7 mV Check: | -22.1 |  |  | Broken. Unoperable. | -7.4 | [8 empty] | pH 7 mV Check:
| Electrolyte replaced Y/N: | Y |  |  |  | N | [8 empty] | Electrolyte replaced Y/N:
| Post mV Check: | -4.3 |  |  |  | N/A | [8 empty] | Post mV Check:
| Depth Calibration: | FAIL |  |  |  | PASS | [8 empty] | Depth Calibration:
| RDO Days: | 366 |  |  |  | 93 | [8 empty] | RDO Days:
| Notes: | Checked early for groundwater sampling. |  |  |  | Checked early after complaint of being slow. | [8 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 5 | SMARTROLL 7 | SMARTROLL 9 | SMARTROLL 10 | SMARTROLL 11 | SMARTROLL 12 | SMARTROLL 13 | AQUATROLL 1 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4
2020-09-02 | Date: |  | 2020-09-02 | 2020-09-02 |  |  | 2020-09-02 | 2020-09-02 | 2020-09-02 | 2020-09-02 | 2020-09-02 | 2020-09-02 | 2020-09-02 | 2020-09-02 | Date:
| pH 7 mV Check: |  | -23.5 | -15.8 | Broken. Unoperable. |  | -17.8 | -21.1 | -20.5 | -16.5 | -5.3 | -19.5 | -13.4 | -18.7 | pH 7 mV Check:
| Electrolyte replaced Y/N: |  | Y | N |  |  | Y | Y | Y | N | N | Y | N | N | Electrolyte replaced Y/N:
| Post mV Check: |  | -5 | N/A |  |  | -1.6 | 1.5 | -6.6 | N/A | N/A | -3.8 | N/A | N/A | Post mV Check:
| Depth Calibration: |  | FAIL | FAIL |  |  | PASS | PASS | FAIL | PASS | PASS | PASS | PASS | PASS | Depth Calibration:
| RDO Days: |  | 456 | 178 |  |  | 178 | 323 | 245 | 128 | ?? | ?? | ?? | ?? | RDO Days:
| Notes: |  | RDO cap replaced at same time. |  |  |  | pH electrolyte replaced as mVs have jumped up quite quickly. | [7 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 5 | SMARTROLL 7 | SMARTROLL 9 | SMARTROLL 10 | SMARTROLL 11 | SMARTROLL 12 | SMARTROLL 13 | AQUATROLL 1 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4
2020-10-01 | Date: | 2020-10-01 | 2020-09-29 | 2020-10-01 | 2020-10-01 | 2020-10-01 | 2020-10-01 | 2020-10-01 | 2020-10-01 | 2020-10-01 | 2020-10-01 | 2020-10-01 | 2020-10-01 | 2020-10-01 | Date:
| pH 7 mV Check: | Not on site | -20 | -22.7 | Broken. Unoperable. | -16.8 | -15.5 | -12.6 | -17.7 | -23.5 | -9.4 | -14.3 | -18.1 | -25.8 | pH 7 mV Check:
| Electrolyte replaced Y/N: |  | Y | Y |  | N | N | N | N | Y | N | N | N | Y | Electrolyte replaced Y/N:
| Post mV Check: |  | -10.1 | -3.1 |  | N/A | N/A | N/A | N/A | -8.9 | N/A | N/A | N/A | -7.9 | Post mV Check:
| Depth Calibration: |  | FAIL | FAIL |  | PASS | PASS | PASS | FAIL | PASS | PASS | PASS | PASS | PASS | Depth Calibration:
| RDO Days: |  | 429 | 149 |  | 58 | 149 | 294 | 217 | 99 | ?? | ?? | ?? | ?? | RDO Days:
| Notes: | [13 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 5 | SMARTROLL 7 | SMARTROLL 9 | SMARTROLL 10 | SMARTROLL 11 | SMARTROLL 12 | SMARTROLL 13 | AQUATROLL 1 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4
2020-10-01 | Date: | 2020-10-29 | 2020-10-29 | 2020-10-29 | 2020-10-29 | 2020-10-29 | 2020-10-29 | 2020-10-29 | 2020-10-29 | 2020-10-29 | 2020-10-29 | 2020-10-29 | 2020-10-29 | 2020-10-29 | Date:
| pH 7 mV Check: | Not on site | -18.6 | Meter not working - removed from fleet. | Broken. Unoperable. | -24.4 | -23.1 | -9.4 | -25.9 | -16.4 | -20.3 | -21.2 | -28.5 | -18.8 | pH 7 mV Check:
| Electrolyte replaced Y/N: |  | N |  |  | Y | Y | N | Y | N | Y | Y | Y | N | Electrolyte replaced Y/N:
| Post mV Check: |  | N/A |  |  | -4.7 | -7.2 | N/A | -10.9 | N/A | -3.2 | -6 | -8.1 | N/A | Post mV Check:
| Depth Calibration: |  | FAIL |  |  | PASS | PASS | PASS | FAIL | PASS | PASS | PASS | PASS | PASS | Depth Calibration:
| RDO Days: |  | 399 |  |  | 30 | ?? | 266 | 188 | 71 | ?? | ?? | ?? | ?? | RDO Days:
| Notes: |  |  |  |  | RDO cap replaced 3/11/2020 | RDO cap replaced 3/11/2020. Has an insitu baro unit now. |  |  | RDO cap replaced 3/11/2020 |  |  |  |  | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 5 | SMARTROLL 7 | SMARTROLL 9 | SMARTROLL 10 | SMARTROLL 11 | SMARTROLL 12 | SMARTROLL 13 | AQUATROLL 1 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4
2020-11-25 | Date: | 2020-11-25 | 2020-11-25 | 2020-11-25 | 2020-11-25 | 2020-11-25 | 2020-11-25 | 2020-11-25 | 2020-11-25 | 2020-11-25 | 2020-11-25 | 2020-11-25 | 2020-11-25 | 2020-11-25 | Date:
| pH 7 mV Check: | -30.3 | -20 | Meter not working - removed from fleet. | Broken. Unoperable. | -12.5 | -12 | -15.9 | -16.3 | -12.2 | -12.4 | -13.1 | -10.1 | -24.8 | pH 7 mV Check:
| Electrolyte replaced Y/N: | Y | Y |  |  | N | N | N | N | N | N | N | N | Y | Electrolyte replaced Y/N:
| Post mV Check: | -7.1 | -7.3 |  |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | -9.3 | Post mV Check:
| Depth Calibration: | FAIL | FAIL |  |  | PASS | PASS | PASS | FAIL | PASS | PASS | PASS | PASS | PASS | Depth Calibration:
| RDO Days: | 276 | 372 |  |  | 343 | ?? | 239 | 161 | 343 | ?? | ?? | ?? | ?? | RDO Days:
| Notes: | [13 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 5 | SMARTROLL 7 | SMARTROLL 9 | SMARTROLL 10 | SMARTROLL 11 | SMARTROLL 12 | SMARTROLL 13 | AQUATROLL 1 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4
2020-12-23 | Date: | 2020-12-23 | 2020-12-23 | 2020-12-23 | 2020-12-23 | 2020-12-23 | 2020-12-23 | 2020-12-23 | 2020-12-23 | 2020-12-23 | 2020-12-23 | 2020-12-23 | 2020-12-23 | 2020-12-23 | Date:
| pH 7 mV Check: | Not on site | -16.8 | Meter not working - removed from fleet. | Broken. Unoperable. | -10.8 | -19.8 | -23 | Not on site | -13.8 | -15.5 | -17.9 | -19.7 | -17.7 | pH 7 mV Check:
| Electrolyte replaced Y/N: |  | N |  |  | N | N | Y |  | N | N | N | N | N | Electrolyte replaced Y/N:
| Post mV Check: |  | N/A |  |  | N/A | N/A | -3.9 |  | N/A | N/A | N/A | N/A | N/A | Post mV Check:
| Depth Calibration: |  | FAIL |  |  | PASS | PASS | PASS |  | PASS | PASS | PASS | PASS | PASS | Depth Calibration:
| RDO Days: |  | 345 |  |  | 315 | ?? | 211 |  | 315 | ?? | ?? | ?? | ?? | RDO Days:
| Notes: | [13 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 5 | SMARTROLL 7 | SMARTROLL 9 | SMARTROLL 10 | SMARTROLL 11 | SMARTROLL 12 | SMARTROLL 13 | AQUATROLL 1 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4
2021-01-08 | Date: | 2021-01-08 | 2021-01-08 | 2021-01-08 | 2021-01-08 | 2021-01-08 | 2021-01-08 | 2021-01-08 | 2021-01-08 | 2021-01-08 | 2021-01-11 | 2021-01-08 | 2021-01-08 | 2021-01-08 | Date:
| pH 7 mV Check: | Not on site | -17.9 | WRITTEN OFF | WRITTEN OFF | -15 | -18.3 | -5.8 | WRITTEN OFF | -15 | not noted | -19.9 | -20.4 | -16.1 | pH 7 mV Check:
| Electrolyte replaced Y/N: |  | Y |  |  | Y | y | N |  | Y | Y | Y | Y | Y | Electrolyte replaced Y/N:
| Post mV Check: |  | -3.9 |  |  | -1.3 | -3.7 | N/A |  | -2.6 | -0.5 | -2.3 | -3 | -4.1 | Post mV Check:
| Depth Calibration: |  | FAIL |  |  | PASS | PASS | PASS |  | PASS | PASS | PASS | PASS | PASS | Depth Calibration:
| RDO Days: |  | 328 |  |  | 299 | ?? | 195 |  | 299 | ?? | ?? | ?? | ?? | RDO Days:
| Notes: | [5 empty] | VuSitu App |  |  |  | pH/ORP probe replaced |  |  |  | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 5 | SMARTROLL 7 | SMARTROLL 9 | SMARTROLL 10 | SMARTROLL 11 | SMARTROLL 12 | SMARTROLL 13 | AQUATROLL 1 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4
2021-02-04 | Date: | 2021-02-03 | 2021-02-03 | 2021-02-03 | 2021-02-03 | 2021-02-03 | 2021-02-03 | 2021-02-03 | 2021-02-03 | 2021-02-03 | 2021-02-03 | 2021-02-04 | 2021-02-03 | 2021-02-03 | Date:
| pH 7 mV Check: | -28.4 | -12.7 | WRITTEN OFF | WRITTEN OFF | -10.9 | -10.8 | -16 | WRITTEN OFF | -12.5 | -12.2 | -11.7 | -15.3 | -17.5 | pH 7 mV Check:
| Electrolyte replaced Y/N: | Y | N |  |  | N | N | N |  | N | N | N | N | N | Electrolyte replaced Y/N:
| Post mV Check: | -5.3 | N/A |  |  | N/A | N/A | N/A |  | N/A | N/A | N/A | N/A | N/A | Post mV Check:
| Depth Calibration: | FAIL | FAIL |  |  | PASS | PASS | PASS |  | PASS | PASS | PASS | PASS | PASS | Depth Calibration:
| RDO Days: | 206 | 302 |  |  | 273 | ?? | 169 |  | 273 | ?? | ?? | ?? | ?? | RDO Days:
| Notes: | [5 empty] | VuSitu App | [7 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 5 | SMARTROLL 7 | SMARTROLL 9 | SMARTROLL 10 | SMARTROLL 11 | SMARTROLL 12 | SMARTROLL 13 | AQUATROLL 1 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4
2021-03-04 | Date: | 2021-02-25 | 2021-03-04 |  |  | 2021-03-04 | 2021-03-04 | 2021-03-04 |  | 2021-03-04 | 2021-03-04 | 2021-03-04 | 2021-03-04 | 2021-03-04 | Date:
| pH 7 mV Check: | -18.2 | -23 | WRITTEN OFF | WRITTEN OFF | -25 | -17.2 | -24.2 | WRITTEN OFF | -19 | -20.5 | -16.8 | -18.9 | -23.9 | pH 7 mV Check:
| Electrolyte replaced Y/N: | Y | Y |  |  | Y | N | Y |  | N | Y | N | N | Y | Electrolyte replaced Y/N:
| Post mV Check: | -6.6 | -4.9 |  |  | -2 | N/A | -4.7 |  | N/A | -4.2 | N/A | N/A | -6.1 | Post mV Check:
| Depth Calibration: | FAIL | FAIL |  |  | PASS | PASS | PASS |  | PASS | PASS | PASS | PASS | PASS | Depth Calibration:
| RDO Days: | 184 | 273 |  |  | 244 | ?? | 140 |  | 244 | ?? | ?? | ?? | ?? | RDO Days:
| Notes: | [13 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 5 | SMARTROLL 7 | SMARTROLL 9 | SMARTROLL 10 | SMARTROLL 11 | SMARTROLL 12 | SMARTROLL 13 | AQUATROLL 1 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4
2021-04-01 | Date: | 2021-04-12 | 2021-03-31 |  |  | 2021-03-31 | 2021-03-31 | 2021-03-31 |  | 2021-03-31 | 2021-03-31 | 2021-03-31 | 2021-03-31 | 2021-03-31 | Date:
| pH 7 mV Check: | -24.5 | Not on site | WRITTEN OFF | WRITTEN OFF | -13.1 | -23.3 | -19.3 | WRITTEN OFF | -28.6 | -13.3 | -29 | -37 | -20.2 | pH 7 mV Check:
| Electrolyte replaced Y/N: | Y |  |  |  | N | Y | Y |  | Y | N | Y | Y | Y | Electrolyte replaced Y/N:
| Post mV Check: | -9.1 |  |  |  | N/A | -3.7 | -7.5 |  | -9.4 | N/A | -7.2 | -7.7 | -10.4 | Post mV Check:
| Depth Calibration: | PASS |  |  |  | PASS | PASS | PASS |  | PASS | PASS | PASS | PASS | PASS | Depth Calibration:
| RDO Days: | 138 |  |  |  | 217 | ?? | 113 |  | 217 | ?? | ?? | ?? | ?? | RDO Days:
| Notes: | [5 empty] | VuSitu App | new RDO 12/4 |  |  | new RDO 12/4 |  |  |  | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 5 | SMARTROLL 7 | SMARTROLL 9 | SMARTROLL 10 | SMARTROLL 11 | SMARTROLL 12 | SMARTROLL 13 | AQUATROLL 1 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4
2021-05-03 | Date: | 2021-05-03 | 2021-05-03 |  |  | 2021-05-03 | 2021-05-03 | 2021-05-03 |  | 2021-05-03 | 2021-05-03 | 2021-05-03 | 2021-05-03 | 2021-05-03 | Date:
| pH 7 mV Check: | Not on site | N/A | WRITTEN OFF | WRITTEN OFF | N/A | N/A | N/A | WRITTEN OFF | N/A | N/A | N/A | N/A | N/A | pH 7 mV Check:
| Electrolyte replaced Y/N: |  | Y |  |  | Y | Y | Y |  | Y | Y | Y | Y | Y | Electrolyte replaced Y/N:
| Post mV Check: |  | -9.4 |  |  | -4.4 | 0.2 | -7.1 |  | -9.9 | -9.1 | -6.7 | -6.4 | -10.9 | Post mV Check:
| Depth Calibration: |  | PASS |  |  | PASS | PASS | PASS |  | PASS | PASS | PASS | PASS | PASS | Depth Calibration:
| RDO Days: |  | VuSitu App |  |  | VuSitu App | VuSitu App |  |  | 217 | VuSitu App | VuSitu App | VuSitu App | VuSitu App | RDO Days:
| Notes: |  | Changed to VuSitu App 03/05/2021 |  |  | Changed to VuSitu App 03/05/2021 | New pH/ORP Probe installed 03/05/2021 | SM3 InSitu baro unit added | [6 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 5 | SMARTROLL 7 | SMARTROLL 9 | SMARTROLL 10 | SMARTROLL 11 | SMARTROLL 12 | SMARTROLL 13 | AQUATROLL 1 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4
2021-06-01 | Date: | 2021-06-01 | 2021-06-01 |  |  | 2021-06-01 | 2021-06-01 | 2021-06-01 |  | 2021-06-01 | 2021-06-01 | 2021-06-01 | 2021-06-01 | 2021-06-01 | Date:
| pH 7 mV Check: | Not on site | -20.5 | WRITTEN OFF | WRITTEN OFF | -12.7 | -8 | -14.6 | WRITTEN OFF | Broken - temperature sensor stopped working in field. | -17.4 | -12.7 | -15.3 | -21.3 | pH 7 mV Check:
| Electrolyte replaced Y/N: |  | Y |  |  | N | N | N |  |  | N | N | N | Y | Electrolyte replaced Y/N:
| Post mV Check: |  | -11.1 |  |  | N/A | N/A | N/A |  |  | N/A | N/A | N/A | -12.6 | Post mV Check:
| Depth Calibration: |  | PASS |  |  | PASS | PASS | PASS |  |  | PASS | PASS | PASS | PASS | Depth Calibration:
| RDO Days: |  | VuSitu App |  |  | VuSitu App | VuSitu App | 315 |  |  | VuSitu App | VuSitu App | VuSitu App | VuSitu App | RDO Days:
| Notes: | [13 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 9 | SMARTROLL 10 | SMARTROLL 11 | SMARTROLL 13 | AQUATROLL 1 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4 | AQUATROLL 5 | AQUATROLL 6 | AQUATROLL 7
2021-07-01 | Date: | 2021-07-07 | 2021-07-07 | 2021-07-07 | 2021-07-07 | 2021-07-07 | 2021-07-07 | 2021-07-07 | 2021-07-07 | 2021-07-07 | 2021-07-07 | 2021-07-07 | 2021-07-07 | 2021-07-07 | Date:
| pH 7 mV Check: | Not on site | -21.9 | -15.4 | -13.9 | -21.7 | Broken - temperature sensor stopped working in field. | -24.3 | -20 | -24.7 | -21 | -10.7 | -11.9 | -11.3 | pH 7 mV Check:
| Electrolyte replaced Y/N: |  | Y | N | N | Y |  | Y | Y | Y | Y | N | N | N | Electrolyte replaced Y/N:
| Post mV Check: |  | -8.8 | N/A | N/A | -4.9 |  | -9 | -7 | -10.4 | -10.7 | N/A | N/A | N/A | Post mV Check:
| Depth Calibration: |  | PASS | PASS | PASS | PASS |  | PASS | PASS | PASS | PASS | PASS | PASS | PASS | Depth Calibration:
| RDO Days: |  | VuSitu App | VuSitu App | VuSitu App | VuSitu App |  | VuSitu App | VuSitu App | VuSitu App | VuSitu App | VuSitu App | VuSitu App | VuSitu App | RDO Days:
| Notes: | [13 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 9 | SMARTROLL 10 | SMARTROLL 11 | AQUATROLL 1 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4 | AQUATROLL 5 | AQUATROLL 6 | AQUATROLL 7 | AQUATROLL 8
2021-08-01 | Date: | 2021-08-13 | 2021-08-03 | 2021-08-03 | 2021-08-03 | 2021-08-13 | 2021-08-13 | 2021-08-03 | 2021-08-13 | 2021-08-13 | 2021-08-13 | 2021-08-03 | At thermofisher due to broken socket | With GW Scientist | Date:
| pH 7 mV Check: | -25.6 | -16.4 | -19.3 | -17.2 | -25.4 | -19.5 | -16 | -21.6 | -22.5 | -24.7 | -19.5 |  |  | pH 7 mV Check:
| Electrolyte replaced Y/N: | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y |  |  | Electrolyte replaced Y/N:
| Post mV Check: | -4.9 | -9.7 | -3.6 | -3.2 | -12.1 | -9.2 | -8.1 | -9.8 | -11.7 | -10.9 | -8.4 |  |  | Post mV Check:
| Depth Calibration: | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS |  |  | Depth Calibration:
| RDO Days: | 429 | VuSitu App | VuSitu App | VuSitu App | VuSitu App | VuSitu App | VuSitu App | VuSitu App | VuSitu App | VuSitu App | VuSitu App | VuSitu App | VuSitu App | RDO Days:
| Notes: | RDO cap replaced. New pH junction | [12 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 9 | SMARTROLL 10 | SMARTROLL 11 | AQUATROLL 1 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4 | AQUATROLL 5 | AQUATROLL 6 | AQUATROLL 7 | AQUATROLL 8
2021-09-01 | Date: | 2021-09-20 | 2021-09-20 | 2021-09-20 | 2021-09-20 | 2021-09-20 | 2021-08-13 | 2021-09-20 | 2021-09-20 | 2021-09-20 | 2021-09-20 | 2021-09-20 | At thermofisher due to broken socket | With GW Scientist | Date:
| pH 7 mV Check: | Not on site - with GW contractor | -22.4 | -19.3 | -13.4 | -31.4 | -19.6 | Not on site - with Drone Tech | N/A | -26.4 | -24.7 | -17 |  |  | pH 7 mV Check:
| Electrolyte replaced Y/N: |  | Y | Y | N | Y | Y |  | Y | Y | Y | N |  |  | Electrolyte replaced Y/N:
| Post mV Check: |  | -6.2 | -3.6 | N/A | -18.3 | -12.2 |  | -13.4 | -19.1 | -10.9 | N/A |  |  | Post mV Check:
| Depth Calibration: |  | PASS | PASS | PASS | PASS | PASS |  | PASS | PASS | PASS | PASS |  |  | Depth Calibration:
| RDO Days: |  | VuSitu App | VuSitu App | VuSitu App | 204 | VuSitu App |  | VuSitu App | VuSitu App | VuSitu App | VuSitu App | VuSitu App | VuSitu App | RDO Days:
| Notes: |  | New RDO Cap |  |  | Isitu App - Lakes |  |  | New PH/ORP New RDO Cap | New RDO Cap |  |  |  |  | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 9 | SMARTROLL 10 | SMARTROLL 11 | AQUATROLL 1 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4 | AQUATROLL 5 | AQUATROLL 6 | AQUATROLL 7 | AQUATROLL 8
2021-10-01 | Date: | 2021-11-01 | 2021-11-01 | 2021-11-01 | 2021-11-01 | 2021-11-01 | 2021-11-01 | 2021-10-26 | 2021-11-01 | 2021-11-01 | 2021-11-01 | 2021-11-01 | WRITTEN OFF under warranty replaced by AQ9 | With GW Scientist | Date:
| pH 7 mV Check: | new probe | new probe | new probe | -19.4 | Damage to pH/ORP Socket. Sent for repair. NCR to be created | n/a | new probe | n/a | -19 | -24.5 | n/a |  |  | pH 7 mV Check:
| Electrolyte replaced Y/N: | Y | Y | Y | N |  | Y | Y | Y | Y | Y | Y |  |  | Electrolyte replaced Y/N:
| Post mV Check: | -6.9 | -3.8 | -5.3 | N/A |  | -10.5 | -4.6 | -10.5 | -6.3 | -10.2 | -5.5 |  |  | Post mV Check:
| Depth Calibration: | FAIL - used Isitu App | PASS | PASS | PASS |  | PASS | PASS | PASS | PASS | PASS | PASS |  |  | Depth Calibration:
| RDO Days: | 377 | VuSitu App | VuSitu App | VuSitu App |  | VuSitu App | VuSitu App | VuSitu App | VuSitu App | VuSitu App | VuSitu App | VuSitu App | VuSitu App | RDO Days:
| Notes: | New PH/ORP | New PH/ORP | New PH/ORP New RDO Cap | New RDO Cap |  |  | New PH/ORP | [6 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 9 | SMARTROLL 10 | SMARTROLL 11 | AQUATROLL 1 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4 | AQUATROLL 5 | AQUATROLL 6 | AQUATROLL 8 | AQUATROLL 9
2021-11-01 | Date: | 2021-11-25 | 2021-11-25 | 2021-11-25 | 2021-11-25 | 2021-11-25 | 2021-11-25 | 2021-11-25 | 2021-11-25 | 2021-11-25 | 2021-11-25 | 2021-11-25 | With GW Scientist | YET TO BE INITIALISED | Date:
| pH 7 mV Check: | Not on site - with GW contractor FOR SAMPLING | -10.5 | -14.5 | -26.2 | WRITTEN OFF | -13.1 | -19.4 | -13.3 | -15.2 | -12.8 | -12.5 |  |  | pH 7 mV Check:
| Electrolyte replaced Y/N: |  | N | N | Y |  | N | N | N | N | N | N |  |  | Electrolyte replaced Y/N:
| Post mV Check: |  | N/A | N/A | -5 |  | N/A | N/A | N/A | N/A | N/A | N/A |  |  | Post mV Check:
| Depth Calibration: |  | PASS | PASS | PASS |  | PASS | PASS | PASS | PASS | PASS | PASS |  |  | Depth Calibration:
| RDO Days: |  | VuSitu App | VuSitu App | VuSitu App |  | VuSitu App | VuSitu App | VuSitu App | VuSitu App | VuSitu App | VuSitu App | VuSitu App | VuSitu App | RDO Days:
| Notes: | [13 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 9 | SMARTROLL 10 | SMARTROLL 11 | AQUATROLL 1 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4 | AQUATROLL 5 | AQUATROLL 6 | AQUATROLL 8 | AQUATROLL 9
2021-12-01 | Date: | 2021-12-23 | 2021-12-23 | 2021-12-23 | 2021-12-23 | 2021-12-23 | 2021-12-23 | 2021-12-23 | 2021-12-23 | 2021-12-23 | 2021-12-23 | 2021-12-23 | With GW Scientist | YET TO BE INITIALISED | Date:
| pH 7 mV Check: | Not on site - with GW contractor FOR SAMPLING | -12.6 | -18 | Not on site | WRITTEN OFF | -21.2 | -25.7 | -20.8 | -23.8 | -18.3 | -19.2 |  |  | pH 7 mV Check:
| Electrolyte replaced Y/N: |  | N | Y |  |  | Y | Y | Y | Y | Y | Y |  |  | Electrolyte replaced Y/N:
| Post mV Check: |  | N/A | -2 |  |  | -7.9 | -3.2 | -4 | -5.5 | -3.2 | -4.4 |  |  | Post mV Check:
| Depth Calibration: |  | FAIL | PASS |  |  | N/A | N/A | N/A | N/A | N/A | N/A |  |  | Depth Calibration:
| RDO Days: |  | 271 | 313 |  |  | PASS | PASS | PASS | PASS | PASS | PASS |  |  | RDO Days:
| Notes: |  | SM Baro Unit 11 | SM Baro Unit 11 |  |  | VuSitu App | VuSitu App | VuSitu App | VuSitu App | VuSitu App | VuSitu App |  |  | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 9 | SMARTROLL 10 | SMARTROLL 11 | AQUATROLL 1 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4 | AQUATROLL 5 | AQUATROLL 6 | AQUATROLL 8 | AQUATROLL 9
2022-01-01 | Date: | 2022-01-10 | 2022-01-10 | 2022-01-10 | 2022-01-10 | 2022-01-10 | 2022-01-10 | 2022-01-10 | 2022-01-10 | 2022-01-10 | 2022-01-10 | 2022-01-10 | With GW Scientist | YET TO BE INITIALISED | Date:
| pH 7 mV Check: | Not on site - with GW contractor FOR SAMPLING | -18.7 | -10 | WRITTEN OFF | WRITTEN OFF | -7.9 | -10.1 | -9.2 | -11.8 | -8.4 | -8.6 |  |  | pH 7 mV Check:
| Electrolyte replaced Y/N: |  | Y | N |  |  | N | N | N | N | N | N |  |  | Electrolyte replaced Y/N:
| Post mV Check: |  | -2.2 | N/A |  |  | N/A | N/A | N/A | N/A | N/A | N/A |  |  | Post mV Check:
| Depth Calibration: |  | PASS | PASS |  |  | PASS | PASS | PASS | PASS | PASS | PASS |  |  | Depth Calibration:
| RDO Days: |  | VuSitu App | VuSitu App |  |  | VuSitu App | VuSitu App | VuSitu App | VuSitu App | VuSitu App | VuSitu App |  |  | RDO Days:
| Notes: | [13 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 9 | AQUATROLL 1 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4 | AQUATROLL 5 | AQUATROLL 6 | AQUATROLL 8 | AQUATROLL 9
2022-02-01 | Date: | 2022-02-01 | 2022-01-31 | 2022-01-31 | 2022-01-31 | 2022-01-31 | 2022-01-31 | 2022-01-31 | 2022-01-31 | 2022-01-31 | With GW Scientist | 2022-01-31 |  |  | Date:
| pH 7 mV Check: | -11.7 | -5.8 | 1.3 | -10.5 | -20.8 | -18.5 | -23 | -16 | -19.5 |  | -4.3 |  |  | pH 7 mV Check:
| Electrolyte replaced Y/N: | N | N | N | N | Y | Y | Y | N | Y |  | N |  |  | Electrolyte replaced Y/N:
| Post mV Check: | N/A | N/A | N/A | N/A | -4.4 | -0.7 | -6.3 | N/A | -7.1 |  | N/A |  |  | Post mV Check:
| Depth Calibration: | FAIL | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS |  | PASS |  |  | Depth Calibration:
| RDO Days: | +284 | VuSitu App (+232) | VuSitu App | VuSitu App | VuSitu App | VuSitu App | VuSitu App | VuSitu App | VuSitu App |  | VuSitu App |  |  | RDO Days:
| Notes: | dropped off with contractor 01/02/22 |  | New PH/ORP | [10 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 9 | AQUATROLL 1 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4 | AQUATROLL 5 | AQUATROLL 6 | AQUATROLL 8 | AQUATROLL 9
2022-03-01 | Date: | 2022-03-30 | 2022-03-29 | 2022-03-29 | 2022-03-29 | 2022-03-29 | 2022-03-29 | 2022-03-29 | 2022-03-29 | 2022-03-29 | With GW Scientist | 2022-03-29 |  |  | Date:
| pH 7 mV Check: | -20.4 | -21.4 | -16.9 | -25.8 | -25.5 | -23.8 | -18.8 | -27.7 | -19.6 |  | -22.7 |  |  | pH 7 mV Check:
| Electrolyte replaced Y/N: | Y | Y | N | Y | Y | Y | Y | Y | Y |  | Y |  |  | Electrolyte replaced Y/N:
| Post mV Check: | -5.2 | -5.4 | N/A | -0.4 | -3.6 | -0.9 | -4.4 | -4.6 | -5.2 |  | -1 |  |  | Post mV Check:
| Depth Calibration: | FAIL | FAIL | PASS | PASS | PASS | PASS | PASS | PASS | PASS |  | PASS |  |  | Depth Calibration:
| RDO Days: | 228 | 175 | 217 | VuSitu App | VuSitu App | VuSitu App | VuSitu App | VuSitu App | VuSitu App |  | VuSitu App |  |  | RDO Days:
| Notes: | Back with Steve | SM Baro Unit 11 | SM Baro Unit 11 | [10 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 9 | AQUATROLL 1 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4 | AQUATROLL 5 | AQUATROLL 6 | AQUATROLL 8 | AQUATROLL 9
2022-05-01 | Date: | 2022-05-02 | 2022-05-02 | 2022-05-02 | 2022-05-02 | 2022-05-02 | 2022-05-02 | 2022-05-02 | 2022-05-02 | 2022-05-02 | With GW Scientist | 2022-05-02 |  |  | Date:
| pH 7 mV Check: | Not on site - with GW contractor FOR SAMPLING | -11.5 | -23.5 | -10.1 | -19.1 | -18.6 | -15.6 | -18.4 | -18.6 |  | -19.5 |  |  | pH 7 mV Check:
| Electrolyte replaced Y/N: |  | N | Y | N | Y | N | N | N | N |  | Y |  |  | Electrolyte replaced Y/N:
| Post mV Check: |  | N/A | -7.6 | N/A | -6.1 | N/A | N/A | N/A | N/A |  | -5.9 |  |  | Post mV Check:
| Depth Calibration: |  | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS |  | PASS |  |  | Depth Calibration:
| RDO Days: |  | VuSitu App | VuSitu App | VuSitu App | VuSitu App | VuSitu App | VuSitu App | VuSitu App | VuSitu App |  | VuSitu App |  |  | RDO Days:
| Notes: | [13 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 9 | AQUATROLL 1 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4 | AQUATROLL 5 | AQUATROLL 6 | AQUATROLL 8 | AQUATROLL 9
2022-05-01 | Date: | 2022-05-25 | 2022-05-25 | 2022-05-25 | 2022-05-25 | 2022-05-25 | 2022-05-25 | 2022-05-25 | 2022-05-25 | 2022-05-25 | With GW Scientist | 2022-05-25 |  |  | Date:
| pH 7 mV Check: | -28 | -18.1 | -8.8 | -14.1 | -16.5 | -26 | -21.2 | -21.6 | -23.1 |  | -13.7 |  |  | pH 7 mV Check:
| Electrolyte replaced Y/N: | Y | N | N | N | N | Y | Y | Y | Y |  | N |  |  | Electrolyte replaced Y/N:
| Post mV Check: | -8.2 | N/A | N/A | N/A | N/A | -6.9 | -7.1 | -8.4 | -9.1 |  | N/A |  |  | Post mV Check:
| Depth Calibration: | FAIL | FAIL | PASS | PASS | PASS | PASS | PASS | PASS | PASS |  | PASS |  |  | Depth Calibration:
| RDO Days: | 171 | 118 | 160 | VuSitu App | VuSitu App | VuSitu App | VuSitu App | VuSitu App | VuSitu App |  | VuSitu App |  |  | RDO Days:
| Notes: | Back with Steve | SM Baro Unit 11 | SM Baro Unit 11 | [10 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 9 | AQUATROLL 1 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4 | AQUATROLL 5 | AQUATROLL 6 | AQUATROLL 8 | AQUATROLL 9
2022-06-01 | Date: | 2022-06-29 | 2022-06-29 | 2022-06-29 | 2022-06-29 | 2022-06-29 | 2022-06-29 | 2022-06-29 | 2022-06-29 | 2022-06-29 | With GW Scientist | 2022-06-29 |  |  | Date:
| pH 7 mV Check: | Not on site - with GW contractor FOR SAMPLING | -23.8 | -14.5 | -23.5 | -26.9 | -18.6 | -16 | -9.7 | -12.4 |  | -20.2 |  |  | pH 7 mV Check:
| Electrolyte replaced Y/N: |  | Y | N | Y | Y | N | N | N | N |  | Y |  |  | Electrolyte replaced Y/N:
| Post mV Check: |  | -5 | N/A | -6.8 | -8.9 | N/A | N/A | N/A | N/A |  | -3.8 |  |  | Post mV Check:
| Depth Calibration: |  | FAIL | PASS | PASS | PASS | PASS | PASS | PASS | PASS |  | PASS |  |  | Depth Calibration:
| RDO Days: |  |  |  | VuSitu App | VuSitu App | VuSitu App | VuSitu App | VuSitu App | VuSitu App |  | VuSitu App |  |  | RDO Days:
| Notes: |  | SM Baro Unit 11 | SM Baro Unit 11 | [10 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 9 | AQUATROLL 1 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4 | AQUATROLL 5 | AQUATROLL 6 | AQUATROLL 8 | AQUATROLL 9
2022-07-01 | Date: | 2022-07-27 | 2022-07-27 | 2022-07-28 | 2022-07-27 | 2022-07-27 | 2022-07-28 | 2022-07-27 | 2022-07-27 | 2022-07-27 | With GW Scientist | 2022-07-27 |  |  | Date:
| pH 7 mV Check: | Not on site - with GW contractor FOR SAMPLING | -12.2 | -22.1 | -11.9 | -17.8 | -27 | -24.2 | -22.8 | -20.4 |  | -15.9 |  |  | pH 7 mV Check:
| Electrolyte replaced Y/N: |  | N | Y | N | N | Y | Y | Y | Y |  | N |  |  | Electrolyte replaced Y/N:
| Post mV Check: |  | N/A | -13.1 | N/A | N/A | -4.9 | -8 | -5.6 | -6 |  | N/A |  |  | Post mV Check:
| Depth Calibration: |  | FAIL | PASS | PASS | PASS | PASS | PASS | PASS | PASS |  | PASS |  |  | Depth Calibration:
| RDO Days: |  | 55 | 96 | VuSitu App | VuSitu App | VuSitu App | VuSitu App | VuSitu App | VuSitu App |  | VuSitu App |  |  | RDO Days:
| Notes: |  | SM Baro Unit 11 | SM Baro Unit 11 | [10 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 9 | AQUATROLL 1 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4 | AQUATROLL 5 | AQUATROLL 6 | AQUATROLL 8 | AQUATROLL 9
2022-08-01 | Date: | 2022-08-24 | 2022-08-25 | Written off | Written off | 2022-08-24 | 2022-08-25 | 2022-08-25 | Off Site | 2022-08-25 | With GW Scientist | 2022-08-25 |  |  | Date:
| pH 7 mV Check: | -35.3 | -18.3 |  |  | -28.7 | -6.3 | -3.5 |  | -14.9 |  | -22.6 |  |  | pH 7 mV Check:
| Electrolyte replaced Y/N: | Y | N |  |  | Y | Y | Y |  | N |  | Y |  |  | Electrolyte replaced Y/N:
| Post mV Check: | -8.7 | N/A |  |  | -7.9 | N/A | N/A |  | N/A |  | -3.3 |  |  | Post mV Check:
| Depth Calibration: | FAIL | FAIL |  |  | PASS | PASS | PASS |  | PASS |  | PASS |  |  | Depth Calibration:
| RDO Days: | 321 | 26 |  |  | VuSitu App | VuSitu App | VuSitu App |  | VuSitu App |  | VuSitu App |  |  | RDO Days:
| Notes: | DO cap from AQ1 | Issue with DO cap replacement |  |  | DO cap replaced | DO cap replaced + pH probe | DO cap replaced + pH probe | [6 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 9 | AQUATROLL 1 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4 | AQUATROLL 5 | AQUATROLL 6 | AQUATROLL 8 | AQUATROLL 9
2022-09-01 | Date: | Off Site | Will not accept new DO caps - not in use | Written off | Written off | 2022-09-29 | 2022-09-29 | 2022-09-29 | 2022-09-29 | 2022-09-29 | 2022-10-07 | 2022-09-30 |  |  | Date:
| pH 7 mV Check: |  |  |  |  | -19.6 | -6.4 | -23.4 | -22.5 | -22.3 | n/a | -16.4 |  |  | pH 7 mV Check:
| Electrolyte replaced Y/N: |  |  |  |  | Y | N | Y | Y | Y | Y | Y |  |  | Electrolyte replaced Y/N:
| Post mV Check: |  |  |  |  | -8 | N/A | -2.9 | -5.9 | -5.8 | 0.5 | -6.9 |  |  | Post mV Check:
| Depth Calibration: |  |  |  |  | PASS | PASS | PASS | PASS | PASS | PASS | PASS |  |  | Depth Calibration:
| RDO Days: |  |  |  |  | VuSitu App | VuSitu App | VuSitu App | VuSitu App | VuSitu App | VuSitu App | VuSitu App |  |  | RDO Days:
| Notes: |  |  |  |  | Temp val 20221007 |  | Broken junction replaced. |  |  | temp val. New pH probe. New RDO cap | previous EOD was 7.20pH |  |  | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | SMARTROLL 9 | AQUATROLL 1 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4 | AQUATROLL 5 | AQUATROLL 6 | AQUATROLL 8 | AQUATROLL 9
2022-10-01 | Date: | Off Site | Will not accept new DO caps - not in use | Written off | Written off | 2022-11-02 | 2022-11-02 | 2022-11-02 | 2022-11-02 | 2022-11-02 | Out of service | 2022-11-02 |  |  | Date:
| pH 7 mV Check: |  |  |  |  | -4.9 | -17.2 | -13.7 | -15.8 | -19.3 |  | -15 |  |  | pH 7 mV Check:
| Electrolyte replaced Y/N: |  |  |  |  | N | N | N | N | N |  | N |  |  | Electrolyte replaced Y/N:
| Post mV Check: |  |  |  |  | N/A | N/A | N/A | N/A | N/A |  | N/A |  |  | Post mV Check:
| Depth Calibration: |  |  |  |  | PASS | PASS | PASS | PASS | PASS |  | PASS |  |  | Depth Calibration:
| RDO Days: |  |  |  |  | VuSitu App | VuSitu App | VuSitu App | VuSitu App | VuSitu App |  | VuSitu App |  |  | RDO Days:
| Notes: | [13 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4 | AQUATROLL 5 | AQUATROLL 6 | AQUATROLL 8 | AQUATROLL 9 | AQUATROLL 10 | AQUATROLL 11 | AQUATROLL 12 | AQUATROLL 13
2022-11-01 | Date: | 2022-11-29 | Will not accept new DO caps - not in use | 2022-11-29 | 2022-12-05 | 2022-11-29 | 2022-12-05 | 2022-11-29 | 2022-11-29 | 2022-11-29 | 2022-11-29 | 2022-11-29 | 2022-12-07 | 2022-11-29 | Date:
| pH 7 mV Check: | -34.2 |  | -11.8 | -14.5 | -7.1 | -25.3 | -28.2 | -9.1 | -24.4 | with env educator | -0.3 | -4.1 | -0.3 | pH 7 mV Check:
| Electrolyte replaced Y/N: | Y |  | N | N | N | Y | Y | N | Y |  | N | N | N | Electrolyte replaced Y/N:
| Post mV Check: | -5.7 |  | N/A | N/A | N/A | -6.8 | -7.9 | N/A | -5.9 |  | N/A | N/A | N/A | Post mV Check:
| Depth Calibration: | FAIL |  | PASS | PASS | PASS | PASS | PASS | PASS | PASS |  | PASS | PASS | PASS | Depth Calibration:
| RDO Days: | 224 |  | VuSitu App | VuSitu App | VuSitu App | VuSitu App | VuSitu App | VuSitu App | VuSitu App |  | VuSitu App | VuSitu App | VuSitu App | RDO Days:
| Notes: | New batteries | [9 empty] | NEW AQ | NEW AQ | NEW AQ | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4 | AQUATROLL 5 | AQUATROLL 6 | AQUATROLL 8 | AQUATROLL 9 | AQUATROLL 10 | AQUATROLL 11 | AQUATROLL 12 | AQUATROLL 13
2023-01-01 | Date: | 2023-01-09 | 2023-01-09 | 2023-01-09 | 2023-01-09 | 2023-01-09 | 2023-01-09 | 2023-01-09 | 2023-01-09 | 2023-01-09 | 2023-01-09 | 2023-01-09 | 2023-01-09 | 2023-01-09 | Date:
| pH 7 mV Check: | Off Site | Will not accept new DO caps - not in use | -21.1 | -21.9 | -17.5 | -13.3 | -18 | -24.5 | -16.2 | with env educator | -15.8 | -13 | -16.3 | pH 7 mV Check:
| Electrolyte replaced Y/N: |  |  | Y | Y | Y | N | Y | Y | Y |  | Y | N | Y | Electrolyte replaced Y/N:
| Post mV Check: |  |  | -4.4 | -6.1 | -5.2 | N/A | -5.7 | -2.4 | -4.6 |  | -4.7 | N/A | -5.3 | Post mV Check:
| Depth Calibration: |  |  | PASS | PASS | PASS | PASS | PASS | PASS | PASS |  | PASS | PASS | PASS | Depth Calibration:
| RDO Calibration: |  |  | PASS | PASS | PASS | PASS | PASS | PASS | PASS |  | PASS | PASS | PASS | RDO Days:
| Date/time check: |  |  | Y | Y | Y | Y | Y | Y | Y |  | Y | Y | Y
| Notes: | [5 empty] | needs temp check |  | GW only | [5 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4 | AQUATROLL 5 | AQUATROLL 6 | AQUATROLL 8 | AQUATROLL 9 | AQUATROLL 10 | AQUATROLL 11 | AQUATROLL 12 | AQUATROLL 13
2023-02-01 | Date: | 2023-02-01 | 2023-02-01 | 2023-02-01 | 2023-02-01 | 2023-02-01 | 2023-02-01 | 2023-02-01 | 2023-02-01 | 2023-02-01 | 2023-02-01 | 2023-02-01 | 2023-02-01 | 2023-02-01 | Date:
| pH 7 mV Check: | Off Site | Will not accept new DO caps - not in use | -13.1 | -15 | -15 | -13.3 | -19.7 | Not checked due for use in March 2023 | -10.1 | with env educator | -15.8 | -15.2 | -22.5 | pH 7 mV Check:
| Electrolyte replaced Y/N: |  |  | N | N | N | N | Y |  | N |  | N | N | Y | Electrolyte replaced Y/N:
| Post mV Check: |  |  | N/A | N/A | N/A | N/A | -8.8 |  | N/A |  | N/A | N/A | -5.8 | Post mV Check:
| Depth Calibration: |  |  | PASS | PASS | PASS | PASS | PASS |  | PASS |  | PASS | PASS | PASS | Depth Calibration:
| Date/time check: |  |  | N | N | N | N | N |  | N |  | N | N | N
| Notes: | [7 empty] | GW only | [5 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4 | AQUATROLL 5 | AQUATROLL 6 | AQUATROLL 8 | AQUATROLL 9 | AQUATROLL 10 | AQUATROLL 11 | AQUATROLL 12 | AQUATROLL 13
2023-03-01 | Date: | 2023-03-01 | 2023-03-01 | 2023-03-01 | 2023-03-01 | 2023-03-01 | 2023-03-01 | 2023-03-01 | 2023-03-01 | 2023-03-01 | 2023-03-01 | 2023-03-01 | 2023-03-01 | 2023-03-01 | Date:
| pH 7 mV Check: | -32.4 | Will not accept new DO caps - not in use | -21.1 | -14.6 | -22.8 | -15.5 | -18.6 | -21.5 | -19.9 | with env educator | Off Site | -14 | -23.8 | pH 7 mV Check:
| Electrolyte replaced Y/N: | Y |  | Y | N | Y | N | N | Y | N |  |  | N | Y | Electrolyte replaced Y/N:
| Post mV Check: | -3.7 |  | -6.1 | N/A | -6 | N/A | N/A | 0.2 | N/A |  |  | N/A | -6.6 | Post mV Check:
| Depth Calibration: | FAIL |  | PASS | PASS | PASS | PASS | PASS | PASS | PASS |  |  | PASS | PASS | Depth Calibration:
| Date/time check: | Y |  | Y | Y | Y | Y | Y | Y | Y |  |  | Y | Y
| Notes: | RDO 132 days | [12 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4 | AQUATROLL 5 | AQUATROLL 6 | AQUATROLL 8 | AQUATROLL 9 | AQUATROLL 10 | AQUATROLL 11 | AQUATROLL 12 | AQUATROLL 13
2023-04-06 | Date: | 2023-04-06 | 2023-04-06 | 2023-04-06 | 2023-04-12 | 2023-04-06 | 2023-04-06 | 2023-04-06 | 2023-04-06 | 2023-04-06 | 2023-04-06 | 2023-04-06 | 2023-04-06 | 2023-04-06 | Date:
| pH 7 mV Check: | With Groundwater Sample Team | Will not accept new DO caps - not in use | -20.5 | -29.2 | -22.9 | -22.4 | -31 | With Groundwater Sample Team | -28.8 | with env educator | -31.9 | -26.1 | -17.6 | pH 7 mV Check:
| Electrolyte replaced Y/N: |  |  | Y | Y | Y | Y | Y |  | Y |  | Y | Y | Y | Electrolyte replaced Y/N:
| Post mV Check: |  |  | -8 | -6.4 | -7.8 | -6.7 | -12.5 |  | -6.6 |  | -7.1 | -9.1 | -10.3 | Post mV Check:
| Depth Calibration: |  |  | PASS | PASS | PASS | PASS | PASS |  | PASS |  | PASS | PASS | PASS | Depth Calibration:
| Date/time check: |  |  | Y | Y | Y | Y | Y |  | Y |  | Y | Y | Y
| Notes: |  |  | corrected NZST | corrected NZST | corrected NZST | corrected NZST | corrected NZST |  | corrected NZST |  | corrected NZST | corrected NZST | corrected NZST | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4 | AQUATROLL 5 | AQUATROLL 6 | AQUATROLL 8 | AQUATROLL 9 | AQUATROLL 10 | AQUATROLL 11 | AQUATROLL 12 | AQUATROLL 13
2023-05-01 | Date: | 2023-05-01 | 2023-05-01 | 2023-05-01 | 2023-05-01 | 2023-05-01 | 2023-05-01 | 2023-05-01 | 2023-05-01 | 2023-05-02 | 2023-05-01 | 2023-05-02 | 2023-05-02 | 2023-05-02 | Date:
| pH 7 mV Check: | With Groundwater Sample Team | Will not accept new DO caps - not in use | -14.7 | -13.4 | -13 | -15.1 | -16.6 | With Groundwater Sample Team | -13.3 | with env educator | -13.7 | -14.8 | -17.3 | pH 7 mV Check:
| Electrolyte replaced Y/N: |  |  | N | N | N | N | N |  | N |  | N | N | N | Electrolyte replaced Y/N:
| Post mV Check: |  |  | N/A | N/A | N/A | N/A | N/A |  | N/A |  | N/A | N/A | N/A | Post mV Check:
| Depth Calibration: |  |  | PASS | PASS | PASS | PASS | PASS |  | PASS |  | PASS | PASS | PASS | Depth Calibration:
| Date/time check: |  |  | Y | Y | Y | Y | Y |  | Y |  | Y | Y | Y
| Notes: | [13 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4 | AQUATROLL 5 | AQUATROLL 6 | AQUATROLL 8 | AQUATROLL 9 | AQUATROLL 10 | AQUATROLL 11 | AQUATROLL 12 | AQUATROLL 13
2023-06-01 | Date: | 2023-06-06 | 2023-05-26 | 2023-05-26 | 2023-05-26 | 2023-05-26 | 2023-05-23 | 2023-05-23 | 2023-06-06 | 2023-05-26 | 2023-05-23 | 2023-05-26 | 2023-05-26 | 2023-05-26 | Date:
| pH 7 mV Check: | -29.5 | Will not accept new DO caps - not in use | -28.8 | -24.7 | -21.1 | -5.3 | -4.8 | -39 | -24 |  | -20.2 | Sent to agent | -26.3 | pH 7 mV Check:
| Electrolyte replaced Y/N: | Y |  | Y | Y | Y | Y | Y |  | Y | Y | Y |  | Y | Electrolyte replaced Y/N:
| Post mV Check: | -3.1 |  | -5 | -6.3 | -6.1 | -5.3 | -4.5 | -1.2 | -4.4 | -6.4 | -4.5 |  | -6.8 | Post mV Check:
| Depth Calibration: |  |  | PASS | PASS | PASS | PASS | PASS |  | PASS | PASS | PASS |  | PASS | Depth Calibration:
| Date/time check: |  |  | Y | Y | Y | Y | Y | Y | Y | Y | Y |  | Y
| Notes: | Quick electrolyte replacement |  |  |  |  | New PH/ORP Probe and RDO Cap installed | New PH/ORP Probe and RDO Cap installed | Quick electrolyte replacement |  | New PH/ORP Probe and RDO Cap installed |  |  |  | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4 | AQUATROLL 5 | AQUATROLL 6 | AQUATROLL 8 | AQUATROLL 9 | AQUATROLL 10 | AQUATROLL 11 | AQUATROLL 12 | AQUATROLL 13
2023-07-01 | Date: | 2023-07-07 | 2023-07-07 | 2023-07-07 | 2023-07-07 | 2023-07-07 | 2023-07-07 | 2023-07-07 | 2023-07-07 | 2023-07-07 | 2023-07-07 | 2023-07-07 | 2023-07-07 | 2023-07-07 | Date:
| pH 7 mV Check: | Will not accept new DO caps - not in use | Will not accept new DO caps - not in use | -18.3 | -21.7 | -19.5 | -16.5 | -19 | Not Checked | -18.5 | with env educator | -19 | With agent | -13.2 | pH 7 mV Check:
| Electrolyte replaced Y/N: |  |  | Y | Y | Y | Y | Y |  | Y |  | Y |  | Y | Electrolyte replaced Y/N:
| Post mV Check: |  |  | -8.4 | -10.3 | -8.5 | -6.1 | -8.1 |  | -8.3 |  | -9 |  | -4.7 | Post mV Check:
| Depth Calibration: |  |  | PASS | PASS | PASS | PASS | PASS |  | PASS |  | PASS |  | PASS | Depth Calibration:
| Date/time check: |  |  | Y | Y | Y | Y | Y |  | Y |  | Y |  | Y
| Notes: | [13 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4 | AQUATROLL 5 | AQUATROLL 6 | AQUATROLL 8 | AQUATROLL 9 | AQUATROLL 10 | AQUATROLL 11 | AQUATROLL 12 | AQUATROLL 13
2023-08-01 | Date: | 2023-08-02 | 2023-08-02 | 2023-08-02 | 2023-08-02 | 2023-08-02 | 2023-08-02 | 2023-08-02 | 2023-08-02 | 2023-08-02 | 2023-08-02 | 2023-08-02 | 2023-08-02 | 2023-08-02 | Date:
| pH 7 mV Check: | Will not accept new DO caps - not in use | Will not accept new DO caps - not in use | -13.4 | -11.5 | -15.2 | -12.3 | -13.9 | Not Checked | -10.6 | with env educator | -19 | With agent | -15.1 | pH 7 mV Check:
| Electrolyte replaced Y/N: |  |  | N | N | N | N | N |  | N |  | N |  | N | Electrolyte replaced Y/N:
| Post mV Check: |  |  | N/A | N/A | N/A | N/A | N/A |  | N/A |  | N/A |  | N/A | Post mV Check:
| Depth Calibration: |  |  | PASS | PASS | PASS | PASS | PASS |  | PASS |  | PASS |  | PASS | Depth Calibration:
| Date/time check: |  |  | Y | Y | Y | Y | Y |  | Y |  | Y |  | Y
| Notes: | [13 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4 | AQUATROLL 5 | AQUATROLL 6 | AQUATROLL 8 | AQUATROLL 9 | AQUATROLL 10 | AQUATROLL 11 | AQUATROLL 12 | AQUATROLL 13
2023-09-01 | Date: |  |  | 2023-09-05 | 2023-09-05 | 2023-09-05 | 2023-09-05 | 2023-09-05 | 2023-09-05 | 2023-09-05 | 2023-09-05 | 2023-09-05 | 2023-09-05 | 2023-09-05 | Date:
| pH 7 mV Check: | Will not accept new DO caps - not in use | Will not accept new DO caps - not in use | Not onsite - Not Checked | -10.5 | -7.8 | -20.1 | -22 | -20.2 | -19.9 | with env educator | -29.2 | With agent | -22.7 | pH 7 mV Check:
| Electrolyte replaced Y/N: |  |  |  | N | N | Y | Y | Y | Y |  | Y |  | Y | Electrolyte replaced Y/N:
| Post mV Check: |  |  |  | N/A | N/A | -3.8 | -7.6 | -1.8 | -4.3 |  | -6.4 |  | -6.4 | Post mV Check:
| Depth Calibration: |  |  |  | PASS | PASS | PASS | PASS | PASS | PASS |  | PASS |  | PASS | Depth Calibration:
| Date/time check: |  |  |  | Y | Y | Y | Y | Y | Y |  | Y |  | Y
| Notes: | [13 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4 | AQUATROLL 5 | AQUATROLL 6 | AQUATROLL 8 | AQUATROLL 9 | AQUATROLL 10 | AQUATROLL 11 | AQUATROLL 12 | AQUATROLL 13
2023-10-01 | Date: | 2023-10-02 | 2023-10-02 | 2023-10-02 | 2023-10-02 | 2023-10-02 | 2023-10-02 | 2023-10-02 | 2023-10-02 | 2023-10-02 | 2023-10-02 | 2023-10-02 | 2023-10-02 | 2023-10-02 | Date:
| pH 7 mV Check: | Will not accept new DO caps - not in use | Will not accept new DO caps - not in use | -17.6 | -20.8 | -19 | -14.2 | -13.6 | Not Checked | -12.3 | with env educator | -15.9 | With agent | -17.6 | pH 7 mV Check:
| Electrolyte replaced Y/N: |  |  | Y | Y | Y | Y | Y |  | Y |  | Y |  | Y | Electrolyte replaced Y/N:
| Post mV Check: |  |  | -9.2 | -8 | -8.3 | -4.2 | -6.3 |  | -4.4 |  | -5.8 |  | -7.1 | Post mV Check:
| Depth Calibration: |  |  | PASS | PASS | PASS | PASS | PASS |  | PASS |  | PASS |  | PASS | Depth Calibration:
| Date/time check: |  |  | Y | Y | Y | Y | Y |  | Y |  | Y |  | Y
| Notes: | [13 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4 | AQUATROLL 5 | AQUATROLL 6 | AQUATROLL 8 | AQUATROLL 9 | AQUATROLL 10 | AQUATROLL 11 | AQUATROLL 12 | AQUATROLL 13
2023-11-01 | Date: | 2023-10-27 | 2023-10-27 | 2023-10-27 | 2023-10-27 | 2023-10-27 | 2023-10-27 | 2023-10-27 | 2023-10-27 | 2023-10-27 | 2023-10-27 | 2023-10-27 | 2023-10-27 | 2023-10-27 | Date:
| pH 7 mV Check: | Will not accept new DO caps - not in use | Will not accept new DO caps - not in use | -13.7 | -13.2 | -13.1 | -9.6 | -13 | Not Checked | -14 | with env educator | -14.4 | With agent | -15.8 | pH 7 mV Check:
| Electrolyte replaced Y/N: |  |  | Y | Y | Y | Y | Y |  | Y |  | Y |  | Y | Electrolyte replaced Y/N:
| Post mV Check: |  |  | -7.8 | -5.4 | -5.2 | 1.1 | -4.8 |  | -1.3 |  | -4 |  | -0.4 | Post mV Check:
| Depth Calibration: |  |  | PASS | PASS | PASS | PASS | PASS |  | PASS |  | PASS |  | PASS | Depth Calibration:
| Date/time check: |  |  | Y | Y | Y | Y | Y |  | Y |  | Y |  | Y
| Notes: | [14 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4 | AQUATROLL 5 | AQUATROLL 6 | AQUATROLL 8 | AQUATROLL 9 | AQUATROLL 10 | AQUATROLL 11 | AQUATROLL 12 | AQUATROLL 13 | AQUATROLL 14
2023-11-01 | Date: | 2023-11-29 | 2023-11-29 | 2023-11-29 | 2023-11-29 | 2023-11-29 | 2023-12-04 | 2023-11-29 | 2023-11-29 | 2023-11-29 | 2023-11-29 | 2023-11-29 | 2023-11-29 | 2023-11-29 | 2023-11-29 | Date:
| pH 7 mV Check: | Will not accept new DO caps - not in use | Will not accept new DO caps - not in use | -17.9 | -16.3 | -17.5 | -13.7 | -18.8 | -7.8 | with service agent | with env educator | -36.6 | With agent | -13.3 | -6.5 | pH 7 mV Check:
| Electrolyte replaced Y/N: |  |  | N | N | N | N | N | N |  |  | Y |  | N | N | Electrolyte replaced Y/N:
| Post mV Check: |  |  | N/A | N/A | N/A | N/A | N/A | N/A |  |  | -7.6 |  | N/A | N/A | Post mV Check:
| Depth Calibration: |  |  | PASS | PASS | PASS | PASS | PASS | PASS |  |  | PASS |  | PASS | PASS | Depth Calibration:
| Date/time check: |  |  | Y | Y | Y | Y | Y | Y |  |  | Y |  | Y | Y
| Notes: | [14 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4 | AQUATROLL 5 | AQUATROLL 6 | AQUATROLL 8 | AQUATROLL 9 | AQUATROLL 10 | AQUATROLL 11 | AQUATROLL 12 | AQUATROLL 13 | AQUATROLL 14
2023-12-01 | Date: | 2024-01-08 | 2024-01-08 | 2024-01-08 | 2024-01-08 | 2024-01-08 | 2024-01-08 | 2024-01-08 | 2024-01-08 | 2024-01-08 | 2024-01-08 | 2024-01-08 | 2024-01-08 | 2024-01-08 | 2024-01-08 | Date:
| pH 7 mV Check: | Will not accept new DO caps - not in use | Will not accept new DO caps - not in use | -28.9 | -28.3 | -27.3 | Not on site | -25.7 | Not on site | with service agent | -50.4 | -31.2 | WRITTEN OFF - REPLACED UNDER WARRANTY WITH AQ15 | Not on site | -21.8 | pH 7 mV Check:
| Electrolyte replaced Y/N: |  |  | Y | Y | Y |  | Y |  |  | Y | Y |  |  | Y | Electrolyte replaced Y/N:
| Post mV Check: |  |  | -2.3 | -4 | -2.9 |  | -2.1 |  |  | -2 | -2.9 |  |  | -1.3 | Post mV Check:
| Depth Calibration: |  |  | PASS | PASS | PASS |  | PASS |  |  | PASS | PASS |  |  | PASS | Depth Calibration:
| Date/time check: |  |  | Y | Y | Y |  | Y |  |  | Y | Y |  |  | Y
| Notes: | [14 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4 | AQUATROLL 5 | AQUATROLL 6 | AQUATROLL 8 | AQUATROLL 9 | AQUATROLL 10 | AQUATROLL 11 | AQUATROLL 13 | AQUATROLL 14 | AQUATROLL 15
2024-01-01 | Date: | 2024-01-30 | 2024-01-30 | 2024-01-30 | 2024-01-30 | 2024-01-30 | 2024-01-30 | 2024-01-30 | 2024-01-30 | 2024-01-30 | 2024-01-30 | 2024-01-30 | 2024-01-30 | 2024-01-30 | 2024-01-30 | Date:
| pH 7 mV Check: | Will not accept new DO caps - not in use | Will not accept new DO caps - not in use | N | -11.3 | -27.3 | N | N | Not on site | with service agent | with env educator | N | N | N | N | pH 7 mV Check:
| Electrolyte replaced Y/N: |  |  | Y | Y | Y | Y | Y |  |  |  | Y | Y | Y | Y | Electrolyte replaced Y/N:
| Post mV Check: |  |  | -9.6 | -10.8 | -2.9 | -5.5 | -8 |  |  |  | -6.3 | -1.1 | -7 | -3 | Post mV Check:
| Depth Calibration: |  |  | PASS | PASS | PASS | PASS | PASS |  |  |  | PASS | PASS | PASS | PASS | Depth Calibration:
| Date/time check: |  |  | Y | Y | Y | Y | Y |  |  |  | Y | Y | Y | Y
| Notes: |  |  |  |  | junction broken- replaced | storage cup broken - replaced | [7 empty] | NEW SONDE | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4 | AQUATROLL 5 | AQUATROLL 6 | AQUATROLL 8 | AQUATROLL 9 | AQUATROLL 10 | AQUATROLL 11 | AQUATROLL 13 | AQUATROLL 14 | AQUATROLL 15
2024-02-01 | Date: | 2024-02-26 | 2024-02-26 | 2024-02-26 | 2024-02-26 | 2024-02-26 | 2024-02-26 | 2024-02-26 | 2024-02-26 | 2024-02-26 | 2024-02-26 | 2024-02-26 | 2024-02-26 | 2024-02-26 | 2024-02-26 | Date:
| pH 7 mV Check: | Will not accept new DO caps - not in use | Will not accept new DO caps - not in use | -19.4 | -17.9 | Not on site | -13.7 | -17.5 | -31.7 | Not on site | with env educator | -20.6 | -12.1 | Not on site | Not on site | pH 7 mV Check:
| Electrolyte replaced Y/N: |  |  | Y | Y |  | N | Y | Y |  |  | Y | N |  |  | Electrolyte replaced Y/N:
| Post mV Check: |  |  | -8.8 | -6.2 |  | N/A | -6.6 | -1.9 |  |  | -6.9 | N/A |  |  | Post mV Check:
| Depth Calibration: |  |  | PASS | PASS |  | PASS | PASS | PASS |  |  | PASS | PASS |  |  | Depth Calibration:
| Date/time check: |  |  | Y | Y |  | Y | Y | Y |  |  | Y | Y
| Notes: | [14 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4 | AQUATROLL 5 | AQUATROLL 6 | AQUATROLL 8 | AQUATROLL 9 | AQUATROLL 10 | AQUATROLL 11 | AQUATROLL 13 | AQUATROLL 14 | AQUATROLL 15
2024-03-01 | Date: | 2024-03-25 | 2024-03-25 | 2024-03-25 | 2024-03-25 | 2024-03-25 | 2024-03-25 | 2024-03-25 | 2024-03-25 | 2024-03-25 | 2024-03-25 | 2024-03-25 | 2024-03-25 | 2024-03-25 | 2024-03-25 | Date:
| pH 7 mV Check: | Will not accept new DO caps - not in use | Will not accept new DO caps - not in use | -15.7 | -14.3 | -24.1 | -20.7 | -15.4 | Not on site | Not on site | with env educator | -16.4 | -19 | -23.2 | -25.4 | pH 7 mV Check:
| Electrolyte replaced Y/N: |  |  | N | N | Y | Y | N |  |  |  | N | N | Y | Y | Electrolyte replaced Y/N:
| Post mV Check: |  |  | N/A | N/A | -7.5 | -2.1 | N/A |  |  |  | N/A | N/A | -5.8 | -6.6 | Post mV Check:
| Depth Calibration: |  |  | PASS | PASS | PASS | PASS | PASS |  |  |  | PASS | PASS | PASS | PASS | Depth Calibration:
| Date/time check: |  |  | Y | Y | Y | Y | Y |  |  |  | Y | Y | Y | Y
| Notes: | [14 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4 | AQUATROLL 5 | AQUATROLL 6 | AQUATROLL 8 | AQUATROLL 9 | AQUATROLL 10 | AQUATROLL 11 | AQUATROLL 13 | AQUATROLL 14 | AQUATROLL 15
2024-04-01 | Date: | 2024-04-23 | 2024-04-23 | 2024-04-23 | 2024-04-23 | 2024-04-23 | 2024-04-23 | 2024-04-23 | 2024-04-23 | 2024-04-23 | 2024-04-23 | 2024-04-23 | 2024-04-23 | 2024-04-24 | 2024-04-24 | Date:
| pH 7 mV Check: | Will not accept new DO caps - not in use | Will not accept new DO caps - not in use | -12.6 | -25 | -15 | -13.2 | -23.4 | Not on site | Not on site | with env educator | -25.3 | -30.1 | -12.1 | -18.3 | pH 7 mV Check:
| Electrolyte replaced Y/N: |  |  | N | Y | N | N | Y |  |  |  | Y | Y | N | N | Electrolyte replaced Y/N:
| Post mV Check: |  |  | N/A | -9 | N/A | N/A | -2.9 |  |  |  | -8.1 | -4.6 | N/A | N/A | Post mV Check:
| Depth Calibration: |  |  | PASS | PASS | PASS | PASS | PASS |  |  |  | PASS | PASS | PASS | PASS | Depth Calibration:
| Date/time check: |  |  | Y | Y | Y | Y | Y |  |  |  | Y | Y | Y | Y
| Notes: | [16 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4 | AQUATROLL 5 | AQUATROLL 6 | AQUATROLL 8 | AQUATROLL 9 | AQUATROLL 10 | AQUATROLL 11 | AQUATROLL 13 | AQUATROLL 14 | AQUATROLL 15 | AQUATROLL 16 | AQUATROLL 17
2024-05-01 | Date: | 2024-05-27 | 2024-05-27 | 2024-05-27 | 2024-05-27 | 2024-05-27 | 2024-05-27 | 2024-05-24 | 2024-05-27 | 2024-05-27 | 2024-05-24 | 2024-05-27 | 2024-05-27 | 2024-05-27 | 2024-05-27 | 2024-05-24 | 2024-05-24 | Date:
| pH 7 mV Check: | Will not accept new DO caps - not in use | Will not accept new DO caps - not in use | -22.7 | -17.6 | -21.5 | -5.7 | -5.4 | -31.9 | Out of service | -5.1 | -14.4 | -11.8 | -19.3 | -26.7 | -1 | -7.3 | pH 7 mV Check:
| Electrolyte replaced Y/N: |  |  | Y | N | Y | N | N | Y |  | N | N | N | Y | Y | N | N | Electrolyte replaced Y/N:
| Post mV Check: |  |  | -11.7 | N/A | -9.8 | N/A | N/A | -7.3 |  | N/A | N/A | N/A | -8.2 | -10.1 | N/A | N/A | Post mV Check:
| Depth Calibration: |  |  | PASS | PASS | PASS | PASS |  | PASS |  |  | PASS | PASS | PASS | PASS |  |  | Depth Calibration:
| Date/time check: |  |  | Y | Y | Y | Y |  | Y |  |  | Y | Y | Y | Y |  |  | Date/time check:
| Notes: |  |  | Firmware upgrade done |  | Firmware upgrade done |  | New pH probe + DO cap installed | Firmware upgrade done |  | New pH probe + DO cap installed |  | Firmware upgrade done |  |  | New pH probe + DO cap installed | New pH probe + DO cap installed | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4 | AQUATROLL 5 | AQUATROLL 6 | AQUATROLL 8 | AQUATROLL 9 | AQUATROLL 10 | AQUATROLL 11 | AQUATROLL 13 | AQUATROLL 14 | AQUATROLL 15 | AQUATROLL 16 | AQUATROLL 17
2024-06-01 | Date: | 2024-06-26 | 2024-06-26 | 2024-06-26 | 2024-06-26 | 2024-07-03 | 2024-06-26 | 2024-07-03 | 2024-06-26 | 2024-06-26 | 2024-06-26 | 2024-07-03 | 2024-06-26 | 2024-06-26 | 2024-06-26 | 2024-06-26 | 2024-06-26 | Date:
| pH 7 mV Check: | Will not accept new DO caps - not in use | Will not accept new DO caps - not in use | -12.9 | -25.3 | -18.4 | -17.2 | -16.5 | Not on site | Out of service | with env educator | -24.6 | -19.7 | -14.2 | -18.4 | -10 | -11.3 | pH 7 mV Check:
| Electrolyte replaced Y/N: |  |  | N | Y | N | N | N |  |  |  | Y | N | N | N | N | N | Electrolyte replaced Y/N:
| Post mV Check: |  |  | N/A | -11.3 | N/A | N/A | N/A |  |  |  | -11.9 | N/A | N/A | N/A | N/A | N/A | Post mV Check:
| Depth Calibration: |  |  | PASS | PASS | PASS | PASS | PASS |  |  |  | PASS | PASS | PASS | PASS | PASS | PASS | Depth Calibration:
| Date/time check: |  |  | Y | Y | Y | Y | Y |  |  |  | Y | Y | Y | Y | Y | Y | Date/time check:
| Notes: | [16 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4 | AQUATROLL 5 | AQUATROLL 6 | AQUATROLL 8 | AQUATROLL 9 | AQUATROLL 10 | AQUATROLL 11 | AQUATROLL 13 | AQUATROLL 14 | AQUATROLL 15 | AQUATROLL 16 | AQUATROLL 17
2024-07-01 | Date: | 2024-07-29 | 2024-07-29 | 2024-07-29 | 2024-07-29 | 2024-07-29 | 2024-07-29 | 2024-07-29 | 2024-07-29 | 2024-07-29 | 2024-07-29 | 2024-07-29 | 2024-07-29 | 2024-07-29 | 2024-07-29 | 2024-07-30 | 2024-07-30 | Date:
| pH 7 mV Check: | Will not accept new DO caps - not in use | Will not accept new DO caps - not in use | -24 | -19.9 | -24 | -22.7 | -24.4 | Not on site | Out of service | with env educator | -17.8 | -27.9 | -21.5 | -28.1 | -16.2 | -21.6 | pH 7 mV Check:
| Electrolyte replaced Y/N: |  |  | Y | Y | Y | Y | Y |  |  |  | Y | Y | Y | Y | Y | Y | Electrolyte replaced Y/N:
| Post mV Check: |  |  | -10.5 | -11.1 | -9.6 | -9.2 | -10.9 |  |  |  | -9.9 | -6 | -6 | -10.7 | -4.6 | -8.9 | Post mV Check:
| Depth Calibration: |  |  | PASS | PASS | PASS | PASS | PASS |  |  |  | PASS | PASS | PASS | PASS | PASS | PASS | Depth Calibration:
| Date/time check: |  |  | Y | Y | Y | Y | Y |  |  |  | Y | Y | Y | Y | Y | Y | Date/time check:
| Notes: | [16 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4 | AQUATROLL 5 | AQUATROLL 6 | AQUATROLL 8 | AQUATROLL 9 | AQUATROLL 10 | AQUATROLL 11 | AQUATROLL 13 | AQUATROLL 14 | AQUATROLL 15 | AQUATROLL 16 | AQUATROLL 17
2024-08-01 | Date: |  |  | 2024-08-30 | 2024-08-30 | 2024-08-30 | 2024-08-30 | Not on site | 2024-08-28 |  |  | 2024-08-30 | 2024-08-30 | 2024-08-30 | 2024-08-30 | 2024-08-30 | 2024-08-30 | Date:
| pH 7 mV Check: | Will not accept new DO caps - not in use | Will not accept new DO caps - not in use | n/a | n/a | n/a | n/a |  | -27 | Out of service | with env educator | n/a | n/a | n/a | n/a | n/a | n/a | pH 7 mV Check:
| Electrolyte replaced Y/N: |  |  | Y | Y | Y | Y |  | Y |  |  | Y | Y | Y | Y | Y | Y | Electrolyte replaced Y/N:
| Post mV Check: |  |  | -4.9 | -4.9 | -0.4 | -6.1 |  | -5.3 |  |  | -6.8 | -1.1 | -5.2 | -6.9 | -1 | -7.9 | Post mV Check:
| Depth Calibration: |  |  | Y | Y | Y | Y |  | PASS |  |  | Y | Y | Y | Y | Y | Y | Depth Calibration:
| Date/time check: |  |  | Y | Y | Y | Y |  | Y |  |  | Y | Y | Y | Y | Y | Y | Date/time check:
| Notes: | [16 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4 | AQUATROLL 5 | AQUATROLL 6 | AQUATROLL 8 | AQUATROLL 9 | AQUATROLL 10 | AQUATROLL 11 | AQUATROLL 13 | AQUATROLL 14 | AQUATROLL 15 | AQUATROLL 16 | AQUATROLL 17
2024-09-01 | Date: | 2024-10-07 | 2024-10-07 | 2024-10-07 | 2024-10-07 | 2024-10-07 | 2024-10-07 | 2024-10-07 | 2024-10-07 | 2024-10-07 | 2024-10-07 | 2024-10-07 | 2024-10-07 | 2024-10-07 | 2024-10-07 | 2024-10-07 | 2024-10-07 | Date:
| pH 7 mV Check: | Will not accept new DO caps - not in use | Will not accept new DO caps - not in use | -15.9 | -11.8 | -17.4 | -17.4 | -14.1 | Not on site | Out of service | with env educator | -18.8 | -10.3 | -19 | -21 | -14.1 | -18.5 | pH 7 mV Check:
| Electrolyte replaced Y/N: |  |  | N | N | N | N | N |  |  |  | N | N | N | Y | N | N | Electrolyte replaced Y/N:
| Post mV Check: |  |  | N/A | N/A | N/A | N/A | N/A |  |  |  | N/A | N/A | N/A | -9.4 | N/A | N/A | Post mV Check:
| Depth Calibration: |  |  | PASS | PASS | PASS | PASS | PASS |  |  |  | PASS | PASS | PASS | PASS | PASS | PASS | Depth Calibration:
| Date/time check: |  |  | Y | Y | Y | Y | Y |  |  |  | Y | Y | Y | Y | Y | Y | Date/time check:
| Notes: | [16 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4 | AQUATROLL 5 | AQUATROLL 6 | AQUATROLL 8 | AQUATROLL 9 | AQUATROLL 10 | AQUATROLL 11 | AQUATROLL 13 | AQUATROLL 14 | AQUATROLL 15 | AQUATROLL 16 | AQUATROLL 17
2024-10-01 | Date: | 2024-10-30 | 2024-10-30 | 2024-10-30 | 2024-10-30 | 2024-10-30 | 2024-10-30 | 2024-10-30 | 2024-10-30 | 2024-10-30 | 2024-10-30 | 2024-10-30 | 2024-10-30 | 2024-10-30 | 2024-10-30 | 2024-10-30 | 2024-10-30 | Date:
| pH 7 mV Check: | Will not accept new DO caps - not in use | Will not accept new DO caps - not in use | -19.9 | -17.7 | -25.5 | -22.9 | -19.6 | Not on site | Out of service | with env educator | -25 | -15.3 | -20.5 | -18.3 | -18.2 | -23.2 | pH 7 mV Check:
| Electrolyte replaced Y/N: |  |  | Y | N | Y | Y | Y |  |  |  | Y | N | Y | Y | N | Y | Electrolyte replaced Y/N:
| Post mV Check: |  |  | -2.7 | N/A | -10.7 | -10.1 | -11 |  |  |  | -9.6 | N/A | -7.1 | -10.4 | N/A | -8.7 | Post mV Check:
| Depth Calibration: |  |  | PASS | PASS | PASS | PASS | PASS |  |  |  | PASS | PASS | PASS | PASS | PASS | PASS | Depth Calibration:
| Date/time check: |  |  | Y | Y | Y | Y | Y |  |  |  | Y | Y | Y | Y | Y | Y | Date/time check:
| Notes: | [16 empty] | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4 | AQUATROLL 5 | AQUATROLL 6 | AQUATROLL 8 | AQUATROLL 9 | AQUATROLL 10 | AQUATROLL 11 | AQUATROLL 13 | AQUATROLL 14 | AQUATROLL 15 | AQUATROLL 16 | AQUATROLL 17
NO-24 | Date: | n/a | n/a | 2024-11-28 | 2024-11-28 | 2024-11-28 | 2024-11-28 | 2024-11-28 | 2024-11-25 | n/a | n/a | 2024-11-28 | 2024-11-28 | 2024-11-25 | 2024-11-28 | 2024-11-28 | 2024-11-28 | Date:
| pH 7 mV Check: | Will not accept new DO caps - not in use | Will not accept new DO caps - not in use | n/a | n/a | n/a | n/a | n/a | n/a | Out of service | with env educator | n/a | n/a | n/a | n/a | n/a | n/a | pH 7 mV Check:
| Electrolyte replaced Y/N: |  |  | Y | Y | Y | Y | Y | Y |  |  | Y | Y | Y | Y | Y | Y | Electrolyte replaced Y/N:
| Post mV Check: |  |  | -3.2 | -4.2 | -9.6 | -5.6 | -5.3 | -6.7 |  |  | -4.6 | -3.2 | -2.3 | -5.9 | -0.5 | -2.6 | Post mV Check:
| Depth Calibration: |  |  | PASS | PASS | PASS | PASS | PASS | PASS |  |  | PASS | PASS | PASS | PASS | PASS | PASS | Depth Calibration:
| Date/time check: |  |  | Y | Y | Y | Y | Y | Y |  |  | Y | Y | Y | Y | Y | Y | Date/time check:
| Notes: | [7 empty] | New RDO Cap and pH Probe |  |  | New RDO Cap and pH Probe | New RDO Cap and pH Probe | New RDO Cap and pH Probe |  |  | replaced storage cup | Notes:
|  | SMARTROLL 2 | SMARTROLL 3 | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4 | AQUATROLL 5 | AQUATROLL 6 | AQUATROLL 8 | AQUATROLL 9 | AQUATROLL 10 | AQUATROLL 11 | AQUATROLL 13 | AQUATROLL 14 | AQUATROLL 15 | AQUATROLL 16 | AQUATROLL 17
2024-12-01 | Date: | n/a | n/a | 2025-01-13 | 2025-01-13 | 2025-01-13 | 2025-01-13 | 2025-01-13 | 2025-01-13 | n/a | n/a | 2025-01-13 | 2025-01-13 | 2025-01-13 | 2025-01-13 | 2025-01-13 | 2025-01-13 | Date:
| pH 7 mV Check: | Will not accept new DO caps - not in use | Will not accept new DO caps - not in use | -18.6 | -17.6 | -21.7 | -44.6 | -22.9 | Not on site | Out of service | with env educator | -15 | -27.7 | Not on site | -23.1 | -14.7 | -19.2 | pH 7 mV Check:
| Electrolyte replaced Y/N: |  |  | Y | Y | Y | Y | Y |  |  |  | Y | Y |  | Y | Y | Y | Electrolyte replaced Y/N:
| Post mV Check: |  |  | -3.7 | -2.7 | -7.5 | -9.1 | -3.2 |  |  |  | -1.7 | -6.1 |  | -7.9 | -2.4 | -6.3 | Post mV Check:
| Depth Calibration: |  |  | PASS | PASS | PASS | PASS | PASS |  |  |  | PASS | PASS |  | PASS | PASS | PASS | Depth Calibration:
| Date/time check: |  |  | Y | Y | Y | Y | Y |  |  |  | Y | Y |  | Y | Y | Y | Date/time check:
| Notes: | [16 empty] | Notes:
|  | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4 | AQUATROLL 5 | AQUATROLL 6 | AQUATROLL 8 | AQUATROLL 9 | AQUATROLL 10 | AQUATROLL 11 | AQUATROLL 13 | AQUATROLL 14 | AQUATROLL 15 | AQUATROLL 16 | AQUATROLL 17
2025-01-01 | Date: | 2025-02-03 | 2025-02-03 | 2025-02-03 | 2025-02-03 | 2025-02-03 | 2025-02-03 | 2025-02-03 | 2025-02-03 | 2025-02-03 | 2025-02-03 | 2025-02-03 | 2025-02-03 | 2025-02-03 | 2025-02-03 | Date:
| pH 7 mV Check: | n/a | n/a | n/a | n/a | n/a | Not on site | Out of service | with env educator | n/a | n/a | n/a | n/a | n/a | n/a | pH 7 mV Check:
| Electrolyte replaced Y/N: | Y | Y | Y | Y | Y |  |  |  | Y | Y | Y | Y | Y | Y | Electrolyte replaced Y/N:
| Post mV Check: | -4.5 | -2.3 | -8.2 | -11.9 | -5.3 |  |  |  | -1.6 | -7.8 | -8.3 | -7 | -1.7 | -0.1 | Post mV Check:
| Depth Calibration: | PASS | PASS | PASS | PASS | PASS |  |  |  | PASS | PASS | PASS | PASS | PASS | PASS | Depth Calibration:
| Date/time check: | Y | Y | Y | Y | Y |  |  |  | Y | Y | Y | Y | Y | Y | Date/time check:
| Notes: | [11 empty] | New RDO Cap |  |  | Notes:
|  | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4 | AQUATROLL 5 | AQUATROLL 6 | AQUATROLL 8 | AQUATROLL 9 | AQUATROLL 10 | AQUATROLL 11 | AQUATROLL 13 | AQUATROLL 14 | AQUATROLL 15 | AQUATROLL 16 | AQUATROLL 17
2025-02-01 | Date: | 2025-02-28 | 2025-02-28 | 2025-02-28 | 2025-02-28 | 2025-02-28 | 2025-02-28 | 2025-02-28 | 2025-02-28 | 2025-02-28 | 2025-02-28 | 2025-02-28 | 2025-02-28 | 2025-02-28 | 2025-02-28 | Date:
| pH 7 mV Check: | n/a | n/a | n/a | n/a | n/a | n/a | Out of service | with env educator | Not on site | n/a | n/a | n/a | n/a | n/a | pH 7 mV Check:
| Electrolyte replaced Y/N: | Y | Y | Y | Y | Y | Y |  |  |  | Y | Y | Y | Y | Y | Electrolyte replaced Y/N:
| Post mV Check: | -3.4 | -4.6 | -7.5 | -5.2 | -7.4 | -5.5 |  |  |  | -5.6 | -4.6 | 0.6 | -1.8 | -8.5 | Post mV Check:
| Depth Calibration: | PASS | PASS | PASS | PASS | PASS | PASS |  |  |  | PASS | PASS | PASS | PASS | PASS | Depth Calibration:
| Date/time check: | Y | Y | Y | Y | Y | Y |  |  |  | Y | Y | Y | Y | Y | Date/time check:
| Notes: | [14 empty] | Notes:
|  | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4 | AQUATROLL 5 | AQUATROLL 6 | AQUATROLL 8 | AQUATROLL 9 | AQUATROLL 10 | AQUATROLL 11 | AQUATROLL 13 | AQUATROLL 14 | AQUATROLL 15 | AQUATROLL 16 | AQUATROLL 17
2025-03-01 | Date: | 2025-03-27 | 2025-03-27 | 2025-03-27 | 2025-03-27 | 2025-03-27 | 2025-03-27 | 2025-03-27 | 2025-03-27 | 2025-03-27 | 2025-03-27 | 2025-03-27 | 2025-03-27 | 2025-03-27 | 2025-03-27 | Date:
| pH 7 mV Check: | n/a | n/a | Not on site | n/a | n/a | Not on site | Out of service | with env educator | n/a | n/a | n/a | n/a | n/a | n/a | pH 7 mV Check:
| Electrolyte replaced Y/N: | Y | Y |  | Y | Y |  |  |  | Y | Y | Y | Y | Y | Y | Electrolyte replaced Y/N:
| Post mV Check: | -6.3 | -5.7 |  | -7.7 | -8.4 |  |  |  | -6.7 | -8.3 | -5.9 | -2.5 | -2.9 | -7.4 | Post mV Check:
| Depth Calibration: | PASS | PASS |  | PASS | PASS |  |  |  | PASS | PASS | PASS | PASS | PASS | PASS | Depth Calibration:
| Date/time check: | Y | Y |  | Y | Y |  |  |  | Y | Y | Y | Y | Y | Y | Date/time check:
| Notes: | [14 empty] | Notes:
|  | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4 | AQUATROLL 5 | AQUATROLL 6 | AQUATROLL 8 | AQUATROLL 9 | AQUATROLL 10 | AQUATROLL 11 | AQUATROLL 13 | AQUATROLL 14 | AQUATROLL 15 | AQUATROLL 16 | AQUATROLL 17
2025-04-01 | Date: | 2025-04-24 | 2025-04-24 | 2025-04-24 | 2025-04-24 | 2025-04-24 | 2025-04-24 | 2025-04-24 | 2025-04-24 | 2025-04-24 | 2025-04-24 | 2025-04-24 | 2025-04-24 | 2025-04-24 | 2025-04-24 | Date:
| pH 7 mV Check: | n/a | n/a | n/a | n/a | n/a | Not on site | Out of service | with env educator | n/a | n/a | n/a | n/a | n/a | n/a | pH 7 mV Check:
| Electrolyte replaced Y/N: | Y | Y | Y | Y | Y |  |  |  | Y | Y | Y | Y | Y | Y | Electrolyte replaced Y/N:
| Post mV Check: | -6.5 | -5.4 | -8.2 | -8.6 | -11.5 |  |  |  | -5.9 | -9.4 | -7 | -5.1 | -3.7 | -9.5 | Post mV Check:
| Depth Calibration: | PASS | PASS | PASS | PASS | PASS |  |  |  | PASS | PASS | PASS | PASS | PASS | PASS | Depth Calibration:
| Date/time check: | Y | Y | Y | Y | Y |  |  |  | Y | Y | Y | Y | Y | Y | Date/time check:
| Notes: | [14 empty] | Notes:
|  | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4 | AQUATROLL 5 | AQUATROLL 6 | AQUATROLL 8 | AQUATROLL 10 | AQUATROLL 11 | AQUATROLL 13 | AQUATROLL 14 | AQUATROLL 15 | AQUATROLL 16 | AQUATROLL 17
2025-05-01 | Date: | 2025-05-29 | 2025-05-29 | 2025-05-29 | 2025-05-28 | 2025-05-28 | 2025-05-19 | 2025-05-28 | 2025-05-29 | 2025-05-29 | 2025-05-29 | 2025-05-29 | 2025-05-29 | 2025-05-29 | Date:
| pH 7 mV Check: | n/a | n/a | n/a | n/a | n/a | n/a | n/a | Not on site | n/a | n/a | n/a | n/a | n/a | pH 7 mV Check:
| Electrolyte replaced Y/N: | Y | Y | Y | Y | Y | Y | Y |  | Y | Y | Y | Y | Y | Electrolyte replaced Y/N:
| Post mV Check: | -7.5 | -6.5 | -10.3 | -7.7 | -7.1 | -9.6 | -8.9 |  | -11.2 | -7.6 | -7.8 | -6.4 | -7.7 | Post mV Check:
| Depth Calibration: | PASS | PASS | PASS | PASS | PASS | PASS | PASS |  | PASS | PASS | PASS | PASS | PASS | Depth Calibration:
| Date/time check: | Y | Y | Y | Y | Y | Y | Y |  | Y | Y | Y | Y | Y | Date/time check:
| Notes: |  |  |  | New RDO Cap and pH Probe | New RDO Cap and pH Probe | prepped for Junes SoE | New RDO Cap and pH Probe |  |  |  |  | New RDO Cap and pH Probe - done 28/5/25 | New RDO Cap and pH Probe - done 28/5/25 | Notes:
|  | AQUATROLL 2 | AQUATROLL 3 | AQUATROLL 4 | AQUATROLL 5 | AQUATROLL 6 | AQUATROLL 8 | AQUATROLL 10 | AQUATROLL 11 | AQUATROLL 13 | AQUATROLL 14 | AQUATROLL 15 | AQUATROLL 16 | AQUATROLL 17
2025-06-01 | Date: | 2025-06-27 | 2025-06-27 | 2025-06-27 | 2025-06-27 | 2025-06-27 | 2025-06-27 | 2025-06-27 | 2025-06-27 | 2025-06-27 | 2025-06-27 | 2025-06-27 | 2025-07-09 | 2025-06-27 | Date:
| pH 7 mV Check: | n/a | n/a | n/a | n/a | n/a | Not on site | with env educator | n/a | n/a | n/a | n/a | -17.6 | n/a | pH 7 mV Check:
| Electrolyte replaced Y/N: | Y | Y | Y | Y | Y |  |  | Y | Y | Y | Y | Y | Y | Electrolyte replaced Y/N:
| Post mV Check: | -5.3 | -4.6 | -8.8 | -7.5 | -2.8 |  |  | -5.5 | -8.9 | -5.1 | -4.6 | -13.2 | -4.3 | Post mV Check:
| Depth Calibration: | PASS | PASS | PASS | PASS | PASS |  |  | PASS | PASS | PASS | PASS | PASS | PASS | Depth Calibration:
| Date/time check: | Y | Y | Y | Y | Y |  |  | Y | Y | Y | Y | Y | Y | Date/time check:
| Notes: | [13 empty] | Notes: